## Slide 1: Preparations for Tomographic Background-Oriented Schlieren Measurements in the 11-by 11-Foot Transonic Wind Tunnel
Joshua M. Weisberger, Brett F. Bathel, Stephen B. Jones
NASA Langley Research Center, Hampton, VA 23666
Mark R. Woike and Jonathon D. Ponder
NASA Glenn Research Center, Cleveland, OH, 44135
James T. Heineck and Edward T. Schairer
NASA Ames Research Center, Moffett Field, CA, 94035
Copyright 2020 United States Government as represented by the Administrator of the National Aeronautics and Space Administration. No copyright is claimed in the United States under Title 17, U.S. Code. All Other Rights Reserved. Published by the American Institute of Aeronautics and Astronautics, Inc. with permission.
AIAA Aviation 2020, June 15 – 19

## Slide 2: Outline
Motivation
Experimental Setup
  Tunnel Facility
  Camera Modules
  Data Acquisition System
  Background Material
  Tunnel Mock-Up
Results
  Camera Window Reflections
  Calibration
  LaRC Candle
  LaRC Heat Gun
  ARC Heat Gun
Conclusions / Lessons Learned

## Slide 3: Motivation
Collaborative opportunity to develop tomographic BOS capability for large NASA wind tunnels
Piggy-back on a test at the Ames 11-by 11-Foot Transonic Wind Tunnel
Limited time available during tunnel testing
  Construct a mock-up to get a feel for setup, acquisition, calibration, and data processing before test
Mock-up constructed at Langley in mid-November 2019
Ames testing in early-to-mid December 2019
Coordinated effort among the three centers
  Langley (LaRC) : camera module development
  Glenn (GRC) : data acquisition system development
  Ames (ARC) : camera positions/orientations in tunnel, background material development

## Slide 4: [Experimental Setup] // 11-by 11-Foot Tunnel
Camera placements dictated by available window locations at the tunnel
[Left]
  Tunnel rendering looking downstream
  Proposed camera locations indicated
  ROI offset from centerline of tunnel
[Top Right]
  Porthole windows for cameras 6 and 7
[Bottom Right]
  Rectangular window for camera 8
Baffled slots not permitted to be covered by the background material
  How would this affect the BOS reconstructions?

## Slide 5: [Experimental Setup] // Camera Module Requirements
Camera module requirements
  One camera per porthole/window around the circumference
  Compact system to fit into tight spaces in plenum behind windows
  Full adjustment of rotation and translation for ideal placement/viewing
  Low-speed acquisition (on the order of a few Hz)
  Withstand low pressures and high temperatures in plenum
  Short exposure to ‘freeze’ fluid flow up to Mach 1.4
  Synchronously trigger all cameras remotely
  On-axis lighting to maximize intensity return
  Minimize reflections from tunnel windows
  Include on-board PCB-based LED driver to maximum intensity and minimize pulse width

## Slide 6: [Experimental Setup] // Camera Module Components
Camera module components (x8)
  Machine vision CMOS camera
  25 mm focal length lens
  Beamsplitter cube
  3D-printed housing which holds:
    Aspheric condenser-diffuser lens
    LED and heat sink
  PCB-based LED driver circuit
  Angled neutral density filter and flat black cap
  3D-printed adapter plate
  Circular arc slot with set screw
Translation/rotation of the camera modules along/about the dashed lines

## Slide 7: [Experimental Setup] // Data Acquisition System
Ethernet cables connect cameras to network switch
Network switch connects to laptop for image acquisition
Function generator and camera trigger lines connect to trigger board
  Trigger the LEDs
  Pulse width of LEDs
  Phase delay between camera and LED trigger
Function Generator
Laptop
Cam 1
Cam 2
Cam 3
Cam 4
Cam 5
Cam 6
Cam 7
Cam 8
Network Switch
To LED PCB
Trigger Board

## Slide 8: [Experimental Setup] // Simultaneous Acquisition
LED Power
Acrylic Windows
LED Trigger Lines

## Slide 9: [Experimental Setup] // Tunnel Mock-Up
Goals
  Detail a calibration method
  Determine whether a successful tomographic reconstruction could be obtained with the camera modules in their proposed positions
  Ensure a quick setup and calibration when system is installed at the tunnel
T-slotted aluminum frame for “tunnel” structure
Camera modules and acrylic “windows” placed at their approximate locations based on renderings/schematics provided by ARC

## Slide 10: [Experimental Setup] // Background Material
Langley Mock-Up
  Alumalite backing panels
  3M Scotchlite 7610 high gain retroreflective sheeting
  Flat black spray paint can with enlarged nozzle to create dot pattern
Ames Tunnel
  7610 adhered to vinyl backing
  From PDF, silk-screen printing of 24” x 63” sections
  Cut to 9.75” width for tunnel slats (between baffled slots)

## Slide 11: [Results] // Camera Module Window Reflection
Why
  Find limits of camera incidence angle to eliminate reflections
How
  Flat pieces of ¼” thick acrylic used to simulate the windows of the tunnel
  Camera module placed 8.75” away from acrylic
  Acrylic rotated about its axis from 0o to 10o degrees, in steps of 2o
0o
2o
4o
6o
8o
10o
Results
  Back-reflection of light source eliminated at 10o
  Vignetting → beamsplitter cube aperture is too small for camera lens (25 mm)
  Could use 50 mm lens to solve this issue, but would not get correct FOV of the model and backgrounds
  Actual FOV for the Ames tunnel test was cut down to a square (white dashed box)
8.75”

## Slide 12: [Results] // Calibration Plate
Cameras all located in roughly same plane
Rod fixed to back of calibration plate
Allowed 360o of rotation of plate, with rigid locking at each image location
Initial plate view defines origin/coordinate system of volume
Images at 13 different views for mock-up calibration
Back of calibration plate

## Slide 13: [Results] // Calibration Results
Translation and rotation from calibration plate origin provided in calibration results
  Plot camera locations for LaRC mock-up and Ames tunnel
Ames tunnel calibration camera locations
Langley mock-up calibration camera locations
Camera locations in mock-up very close to final camera locations in tunnel tests
Cameras 1 and 2 at mock-up located on the floor
Tunnel Walls

## Slide 14: [Results] // Langley Candle Results
What
  Single tea-light candle held at
  center of ROI using a piece of 80/20
Why
  Small density object
  Slight obstruction in some frames due to rod
  Nearly axisymmetric reconstruction in laminar region
Results
  2D BOS images colored by pixel displacement
  3D reconstruction projected back to each camera (a sort of synthetic 2D BOS)
  Ten time frames of the reconstruction (acquired at 4 Hz)

## Slide 15: [Results] // Langley Heat Gun Results : Round Nozzle
What
  Heat gun with round nozzle attachment
Why
  Density object similar size to flight vehicle model
  Quasi-axisymmetric turbulent plume
Results
  2D BOS images colored by pixel displacement
  3D reconstruction projected back to each camera (a sort of synthetic 2D BOS)
  Temperature iso-surfaces, and two plane slices of the reconstruction
Takeaway
3D reconstruction qualitatively looks reasonable
Temperatures are reasonable

## Slide 16: [Results] // Langley Heat Gun Results : Flat Nozzle
What
  Heat gun with flat nozzle attachment
Why
  Density object similar size to flight vehicle model
  Plume shape different in all three dimensions
Results
  2D BOS images colored by pixel displacement
  3D reconstruction projected back to each camera (a sort of synthetic 2D BOS)
  Temperature iso-surfaces, and three plane slices of the reconstruction
Takeaway
Fully 3D object able to be reconstructed successfully
Qualitatively looks reasonable
Temperatures are reasonable

## Slide 17: [Results] // Ames Heat Gun Results : Round Nozzle
What
  Heat gun with round nozzle attachment
Why
  Density object similar size to flight vehicle model
  Quasi-axisymmetric turbulent plume
Results
  2D BOS images colored by pixel displacement
  Reconstruction of the projected 3D reconstruction back to each camera
  Temperature iso-surfaces, and two plane slices of the reconstruction
Takeaway
Slots can be seen in projected reconstruction, but are not noticeable in 3D and slice views
3D result looks qualitatively accurate
Temperatures in plume are reasonable
High level of confidence for actual model testing

## Slide 18: Conclusions / Lessons Learned
Did we accomplish our goals?
  Detail a calibration method
  Successful tomographic reconstruction of a small BOS object with camera modules in proposed tunnel locations
  Quick setup, calibration, acquisition, and teardown at the tunnel to adhere to time allotment
Additional acknowledgments:
  NASA Transformation Tools and Technologies (TTT)
  NASA Aerodynamic Evaluation and Test Capabilities (AETC)
  NASA Commercial Supersonic Transport (CST)


