## Slide 1: Electromagnetic Thermocouple (TC) Anomaly Mechanisms During Atmospheric Entry
Magnus Haw
Sensors and Instrumentation Webcast
April 13, 2021

## Slide 2: What is a Thermocouple?
In metals, electron mobility changes as a function of temperature
Therefore, thermal gradients produce voltage gradients in metals
uV/K
Q: So why is a junction needed?
A: Metrology reasons: voltage meters require closed circuits & physical contacts.
Consequently, TCs use two different material wires & measure the tiny temperature-dependent voltage difference between them

## Slide 3: Why are TC anomalies important to understand?
We rely on TCs for nearly all high temperature measurements
TC high temperature measurements form the basis for many forms of simulation validation (material, aerothermal, radiative)
Important to know if in-flight safety thresholds are being exceeded
Galileo, Shuttle, Orion

## Slide 4: Three Mechanisms
Surface voltage coupling
Hall effect current
RF rectification

## Slide 5: #1. Surface voltage coupling
Short story: surface charge/voltages couple to near-surface bare TCs
Long story:
  Turbulence and pyrolysis gases enhance boundary layer electron density
  Non-parallel electron temperature and pressure gradients generate electric fields and currents in boundary layer & charge heatshield surface
  Non-uniform surface charge creates non-uniform voltages across heatshield
  Surface voltage couples to embedded bare TCs

## Slide 6: #1. Surface voltage coupling: Equations
B-field from generalized Ohm’s Law:
Assume EM fields are quasi-static, i.e., dB/dt = 0, dE/dt = 0
Solve for B, then calculate E

## Slide 7
take divergence of Ohms Law:
#1. Surface voltage coupling: Surface charge

## Slide 8
#1. Surface voltage coupling: TC placement

## Slide 9
Need to solve below electrostatic system for in TC geometry
Wire-TPS conductivity ratio varies from 1000/1 to 10000/1
Solution: use 3D finite element simulations
#1. Surface voltage coupling: Modeling

## Slide 10: Numerical Methods
Used Python API for FEniCS platform
Finite element solver (2nd order): Rayleigh-Ritz (variational calculus method)
Variational (weak) formulation of potential equation
System run on full-domain and half domain.
The half domain mesh contained higher resolution in the yz-plane to better resolve the small wire radius.
All tests gave equivalent results for the wire voltage within 1%.
Numerical convergence

## Slide 11: Assuming linear surface potential: 34-75% coupling

## Slide 12: Assuming dipole surface potential: 17-36% coupling

## Slide 13
The conductivity ratio between the TC wire and the sample is strongly dependent on temperature.
  metal conductivity drops at higher temperatures (5x)
  char conductivity increases at higher temperatures (>2x)
  contribution of pyrolysis gases to conductivity is unknown (??)
Estimate: > 10000/1 at low temperatures (300 K)
Estimate: < 1000/1 at high temperatures (1500 K)
#1. Surface voltage coupling: Conductivity ratio

## Slide 14
Atmospheric entry conditions where boundary layer electron density is >> 1017 m3 (~Tgas > 5000 K)
TCs are embedded near the surface (< 0.25’’)
Anomaly is enhanced by turbulence and TPS materials containing elements with low ionization potentials (e.g., Si).
Anomaly is enhanced at higher material temperatures (char conductivity rises, metal conductivity falls)
#1. Surface voltage coupling: Relevance

## Slide 15
Electrically insulate near-surface TC wires
Reduce conductivity of TPS material/pyrolysis products (remove Si)
Get surface voltage distribution by measuring single-ended voltages of each TC contact, then model & correct for perturbations.
#1. Surface voltage coupling: Mitigation

## Slide 16
Short story: Hall effect current driven through TC
Long story:
  Hall effect from Earth’s magnetic field creates an orientation dependent voltage drop across spacecraft frame
  A low impedance or grounded TC terminal allows current to flow through measurement circuit
  Ionized pyrolysis gases short TC to external plasma
  Circuit is closed via current path through external layer of plasma
#2. Hall effect currents

## Slide 17
VxB electric field
B-field
Velocity
Metal frame

## Slide 18
B-field
Velocity
+
+
+
-
-
-
VxB field
Current flows in metal frame due to
VxB field & begins to build up on edges

## Slide 19
B-field
Velocity
+
+
+
-
-
-
VxB field
Edge plasma is stationary relative to
Earth (VxB=0) & sufficiently conductive
to carry return current

## Slide 20
B-field
Velocity
VxB field
Return current flows through edge plasma
Return current eliminates shielding effects
(i.e. no charge buildup on frame)

## Slide 21
B-field
Velocity
Return current flows through edge plasma
total emf =
TC wire
VxB field
L
current anomaly = emf/R
TC experiences VxB emf
and draws anomalous
current

## Slide 22
Experiment with long wire (20 km) deployed from Shuttle with sensors at terminal ball
Produced kilovolts of EMF & melted tether
Same effect is present for all spacecraft but is orders of magnitude smaller since spacecraft size, L << 20 km
#2. Hall effect currents: Tether Experiment

## Slide 23
100 degC / 1.3 mV
Shuttle anomaly current threshold, @80 Ohms: 34 uA
These anomalous currents are typically too small to affect other diagnostics
Very sensitive!
#2. Hall effect currents: TC Sensitivity

## Slide 24
Spacecraft with metal frame
TCs are embedded near the exterior surface (< 0.25’’)
Magnetic field is present (VxB≠0)
External plasma is electrically connected to TC and metal frame
TCs must have at least one low impedance terminal
Anomaly is enhanced by length of metal frame, spacecraft speed, and magnetic field magnitude/orientation.
#2. Hall effect currents: Relevance

## Slide 25
Make both TC terminals high impedance (easy)
Fully insulate TCs/internal metal frame from external plasma (hard)
Use smaller entry vehicles or lower entry speeds
#2. Hall effect currents: Mitigation

## Slide 26
Short story: RF interference can cause DC offsets in TC measurements
Long story:
  The low voltage produced by TCs (uV level) typically passes through an instrumentation amplifier (inamp) before digitization
  Inamps are slow and cannot reject common mode signals above ~20 kHz
  Radio frequency interference (RFI) gets rectified in TC inamp circuits leading to DC offsets
  Asymmetric rectification can occur from any unbalanced circuit elements
#3. RF Rectification

## Slide 27
Entry environment with significant incident RF power (>> mW)
#3. RF Rectification: Relevance

## Slide 28
Filtering
  RC low pass filter
  Common mode choke
  X2Y capacitor filter
Minimize RF pickup
  reduce circuit area
  use coaxial or twisted pair wire
Shielding
#3. RF Rectification: Mitigation

## Slide 29: TC Anomaly Summary
Surface voltages can couple into near-surface bare TC wires
Low-impedance TC terminals are susceptible to Hall currents
RFI can generate DC offsets in TC digitization
Questions?