Hydrogen Peroxide as a Method for Bioburden Reduction in Facilities with Strict Materials Requirements.

A. B. Regberg1 , C.L. Amick2, R. E. Davis3, E.K. Lewis3, F. Mazhari2, J.L. Mitchell1, D.L. Owens2 and, F. M. McCubbin1 1Astromaterials Research and Exploration Science Division, NASA Johnson Space Center, 2101 NASA Parkway, Houston TX, 77058 2 NASA Johnson Space Center/Jacobs-JETS, 2101 NASA Parkway, Houston TX 77058, 3Texas State University, Johnson Space Center 2101 E NASA Pkwy Houston, TX, 77058, USA

The cleanrooms used to curate NASA’s Astromaterials samples are carefully monitored for particulate and inorganic contamination. The clean labs also have a very limited set of acceptable materials and cleaning agents to further minimize the potential for contamination. Labs are cleaned primarily with isopropyl alcohol. Astromaterials samples are handled with tools made of stainless steel (304 or 316), Teflon, or aluminum alloy (6061). Although our current collections are not particularly susceptible to biological alteration or organic contamination, this will not be the case for new collections from the OSIRIS-REx mission, Hayabusa2, and from Mars Sample Return. Therefore, it is necessary to develop and test methods to reduce the bioburden in astromaterials cleanrooms without introducing unwanted contaminants.

We will report on the results of three case studies where 7.5 wt% hydrogen peroxide was prepared from a stock solution of ultrapure 30 wt% hydrogen peroxide (JT Baker) using curation-grade ultrapure water. We followed CDC (Center for Disease Control) guidelines for using hydrogen peroxide as a high level disinfectant. This solution was used to clean a glovebox prior to processing Apollo samples, as well as surfaces in the Antarctic meteorite processing lab and Stardust lab after facilities monitoring indicated an unwanted increase in bioburden. In all three instances, the culturable bioburden was significantly reduced after a 30 min. exposure to the 7.5% hydrogen peroxide solution without a corresponding increase in inorganic or organic contamination. We observed 77 to 100% reductions in the bioburden recovery rate. In one case study, we also performed amplicon DNA sequencing on samples collected from the surfaces before and after cleaning. We observed a significant change in microbial community composition after peroxide cleaning. These results suggest that routine cleaning with hydrogen peroxide could be an effective way to control bioburden in astromaterials cleanrooms and other facilities with strict contamination control requirements.