## Slide 1: Using Digital Image Correlation for Material Testing at Cryogenic Temperatures
Joseph Cochran, University of Utah
Matlock Mennu, Analytical Services & Materials, Inc.
iDICs 2022

## Slide 2: Overview
NASA Langley Research Center (LaRC) Fatigue and Fracture Lab
Metal Alloy Characterization Based on Temperature
Digital Image Correlation (DIC) and Environmental Chamber Setup
Test
Results
Issues During Testing
Observation and Mitigation
https://images.nasa.gov/details-LRC-2013-B701_P-00137

## Slide 3: NASA LaRC Fatigue and Fracture Lab
Several hydraulic and servomechanical load frames available for material and sub-component level testing
Load Frames: 5,000 lbs. to 400,000 lbs.
Cyclic and static loading
Tests are generally non-standard to characterize behavior and/or provide data for test/analysis correlation
DIC is routinely used for nearly every test
Environmental chambers are available for testing at temperatures ranging from -320°F to 600°F

## Slide 4: Environmental Chamber Setup
Environmental chamber mounted to Instron hydraulic test frame for characterizing material properties
Port holes for fitting to temperature control unit and access to instrumentation inside
Hinged doors on both front and back, with retrofitted quartz glass for observing test coupons
How to mount DIC cameras for the tests

## Slide 5: Initial DIC Setup
Camera setup on aluminum structural framing mounted on the hinged door
  Allows access to the chamber
  Maintain offset from test coupon for multiple samples
  No movement/adjustment of cameras
Plastic tent was used to cover the aluminum framing and DIC cameras to prevent condensation on the windows
Dry air hose was routed through this plastic tent directed at the glass window
Setup was duplicated on both front and back side
Lighting: Light-emitting diode (LED) strips were mounted on left and right side of the glass window
Cam 1
Cam 2

## Slide 6: DIC Camera Setup
Cameras: Basler Ace 12 MP*
Image Resolution: 4096 X 3000 pixels2
Lens: 50-mm lens
Aperture: f/5.6
Field of View: 5 in
Image Scale: 750 pixels/in
Stereo-Angle: 22.28°
Stand-Off Distance: 20 in
Image Acquisition Rate: 1 Hz
Cam 1
Cam 2
* Specific vendor and manufacturer names are explicitly mentioned only to accurately describe the test hardware. The use of vendor and manufacturer names does not imply an endorsement by the U.S. Government nor does it imply that the specified equipment is the best available.

## Slide 7: Calibration
Calibration was performed one side at a time
  The side most frequented for removing and replacing test coupons was calibrated last
  Calibration grid used: 5 mm
  Software used: VIC-3D
Calibration was done at room temperature before starting the temperature cycle and testing

## Slide 8: Testing Round Notch Coupons
Quasi-static loading (0.02 in/min) until failure
Extensometers were used as a fail-safe if DIC results could not be obtained (mainly during cryogenic) temperatures
  Two extensometers mounted on the coupon with springs
DIC data recorded with load and temperature data from the load frame using Bayonet Neill–Concelman (BNC) cables

## Slide 9: Initial Results
Observed data noise levels
  Larger than expected projection errors (>0.1 pixels)
  Suspected camera vibration due to liquid nitrogen (LN2) inflow
Strain errors
  Observed at cold temperatures and strains greater than 5%
  Errors caused by paint cracking
Stress-strain curve
Cracks in paint at 5%
Back Avg: 0.313 pixel
Front Avg: 0.178 pixel
0
0.2
0.4
0.6
0.8
1
1.2
1.4
0
200
400
600
800
Error (px.)
Time(s)
Projection Error
Back
Front

## Slide 10: Coefficient of Thermal Expansion (CTE) Test
CTE experiments
  Measure strain as a function of temperature
  Three temperature holds
    0°F, -75°F, -150°F
Projection error increased with decreased temperature curve
0
0.5
1
1.5
2
2.5
3
3.5
0
5000
10000
15000
20000
25000
30000
35000
40000
45000
50000
55000
60000
65000
70000
75000
80000
Error (px.)
Projection Error and Temperature
-150
-100
-50
0
50
100
150
0
5000
10000
15000
20000
25000
30000
35000
40000
45000
50000
55000
60000
65000
70000
75000
80000
Temp. (
°
F)
Time (s)
SET POINT

## Slide 11: Return to Room Temperature
Chart: Projection Error and Temperature
Projection error decreases with increased temperature
Isolate temperature effects from camera setup

## Slide 12: Error Sources
Vibration
  Temperature control unit connected to chamber
  External fan used to disperse nitrogen escaping the chamber
Temperature
  Liquid nitrogen introduced in the chamber with copper tubing attached to the top of the chamber
    Temperature gradient (mirage effect)
  Frost build up when testing multiple coupons

## Slide 13: Temperature Control Unit
Produces vibration
  Due to blower used for air conditioning
  Solenoid used for regulating LN2 inlet
External fan
  No perfect seals
  Sometimes used to disperse excess LN2 surrounding test frame

## Slide 14: Frost During Testing
Frost build up on coupon during consecutive tests

## Slide 15: New Setup
Use tripod for mounting cameras
  Clamped to the rigid beam structure
Plastic tent only on the glass window
  Minimize the area for dry air distribution
Additional air hose for each camera
Monitor real time on each side with separate computer systems
Tripod
Clamp

## Slide 16: Results
Projection error reduced with new setup
  Measurement error independent from temperature effects
Limitation to the setup
  The cameras needs to be isolated from all vibration and temperature effects while allowing ease of access to inside of chamber
Chart: Projection Error
Tripod
Clamp

## Slide 17: Summary of Tests Completed
Round, notched tensile tests (extracted from hardware component 1)
  3 different notches for different levels of triaxiality
  3 different temperatures
Flat, notched tensile tests (extracted from hardware component 1)
  4 different notches for different level of triaxiality
  3 different temperatures
Shear at 3 different temperatures (extracted from hardware component 1)
Sub-sized tensile tests at 3 different temperatures (extracted from hardware component 2)
CTE tests on material (extracted from hardware component 1 and 3)
Stress-strain and CTE data
Development of a material model for a critical flight hardware

## Slide 18: Thank you
Questions?