## Slide 1: High Temperature Emissivity Measurement System (HiTEMS)
Meghan Carrico (meghan.carrico@nasa.gov)
Space Environmental Effects, NASA Marshall Space Flight Center (Huntsville, AL)
DISTRIBUTION A. Approved for public release; distribution unlimited.

## Slide 2: High Temperature Emissivity Measurement System (HiTEMS)
The High Temperature Emissivity Measurement System (HiTEMS) measures spectral and total near normal emittance
  Capable of a range of incident angles from 0° to 180°
  Record supplemental data of laser power, blackbody and sample temperature, tolerance, and signal
Wavelength range:
  600°C to 1,500°C: approximately 2 µm to 24.9 µm (Experiment-1)
    KBr beam splitter
  1,600°C+: approximately 0.67 µm to 1.04 µm (Experiment-3)
    Quartz beam splitter
Temperature range: approximately 600°C – 2,600°C
  Depends on sample size and treatment
Vacuum levels: as low as approximately 10-7 Torr
DISTRIBUTION A. Approved for public release; distribution unlimited.

## Slide 3: High Temperature Emissivity Measurement System (HiTEMS)
Utilizes a Mikron M390 Blackbody Calibration Source
FAR 12257-1 Pyrometer
FTIR: Nicolet 6700
Laser: 400 W 1,070 nm Ytterbium fiber laser
  IPG Photonics
Laser uses a beam expander to provide uniform heating
  with beam expander: beam diameter of 5 mm
  without beam expander: beam diameter of 3 mm
DISTRIBUTION A. Approved for public release; distribution unlimited.

## Slide 4: HiTEMS
DISTRIBUTION A. Approved for public release; distribution unlimited.
FTIR
Camera
Pyrometer
Laser
Chamber
Blackbody

## Slide 5: High Temperature Emissivity Measurement System (HiTEMS)
Typical sample size: 3/8’’ x 3/8’’
  Circular samples preferred for uniform heating
  Flexible on sample dimensions, can provide in-house machining
Utilize a path factor to correct for differences between blackbody and sample path
  Hexoloy SA used for calibration
DISTRIBUTION A. Approved for public release; distribution unlimited.

## Slide 6: HiTEMS
DISTRIBUTION A. Approved for public release; distribution unlimited.

## Slide 7: HiTEMS Tungsten Emittance Tests
HiTEMS tungsten emittance measurements were completed as a follow-up to high temperature molybdenum emittance measurements
  Tests were conducted for MSFC ER64 Propulsion Research and Tech Branch
Molybdenum measurements were completed to support the development of new heat transfer correlations to aid in thermal analysis of nuclear rocket fuel elements
During Mo tests:
  Tested bead blasted and shiny Mo samples
  Discovered change in emittance from HiTEMS test #1 to HiTEMS test #2 for bead blasted Mo
Subsequent HiTEMS tests of tungsten explored this observation with a bead blasted and smooth/shiny sample
DISTRIBUTION A. Approved for public release; distribution unlimited.

## Slide 8: Specifications/Outline for Tungsten HiTEMS Tests
ER64-001: Shiny tungsten sample
ER64-002: Bead blasted tungsten sample
ER64-001 test conditions
  Test #1: 851 °C to 1,700 °C
ER64-002 test conditions
  Test #1: 851 °C to 2,601 °C
  Test #2: 849 °C to 2,499 °C
  Test #3: 850 °C to 2,503 °C
ER64-001 sample area: 31.77 mm2
ER64-002 sample area: 30.97 mm2
DISTRIBUTION A. Approved for public release; distribution unlimited.

## Slide 9: Pre-Test Photos
ER64-001
ER64-002
DISTRIBUTION A. Approved for public release; distribution unlimited.

## Slide 10
DISTRIBUTION A. Approved for public release; distribution unlimited.

## Slide 11
DISTRIBUTION A. Approved for public release; distribution unlimited.

## Slide 12
DISTRIBUTION A. Approved for public release; distribution unlimited.

## Slide 13
DISTRIBUTION A. Approved for public release; distribution unlimited.
ER64001 & ER64002 Run 1
850°C to 1,500°C

## Slide 14
DISTRIBUTION A. Approved for public release; distribution unlimited.
ER64001 & ER64002 Run 1
1600°C to 2,600°C

## Slide 15
DISTRIBUTION A. Approved for public release; distribution unlimited.
ER64002 All Runs
850°C to 1,500°C

## Slide 16
DISTRIBUTION A. Approved for public release; distribution unlimited.
ER64002 Run 1
1600°C to 2,600°C

## Slide 17
DISTRIBUTION A. Approved for public release; distribution unlimited.
ER64002 Run 2
1600°C to 2,500°C

## Slide 18
DISTRIBUTION A. Approved for public release; distribution unlimited.
ER64002 Run 3
1600°C to 2,500°C

## Slide 19: Observations
As expected, the average total emittance for the first run of the bead blasted tungsten is more than double the shiny tungsten average total emittance
  The shiny tungsten average total emittance is similar to the second and third run of the bead blasted tungsten average total emittance
The bead blasted tungsten initial average total emittance drops by more than half from the first run to later runs
  The average total emittance for all runs slowly realigns at higher temps
The shiny tungsten spectral emittance is much lower than the spectral emittance from the first run of the bead blasted tungsten
For the lower temperature range, the first run of the bead blasted tungsten spectral emittance is much higher than the following runs
  Except for in regions of noise towards the end of the lower temperature range
Not as much variation in the spectral emittance is seen for the bead blasted tungsten through all runs in the high temperature range although the spectral emittance does tend to decrease with each subsequent run
DISTRIBUTION A. Approved for public release; distribution unlimited.

## Slide 20: Laser Power
Sample Temperature [°C] | ER64-001 Laser Power Test 1 [W] | ER64-002 Laser Power Test 1 [W] | ER64-002 Laser Power Test 2 [W] | ER64-002 Laser Power Test 3 [W]
850 | 29.7 | 23.6 | 9.38 | 26.9
900 | 33 | 26.3 | 10.3 | 30
1,000 | 40.9 | 33 | 13.3 | 36.3
1,100 | 49.6 | 41.7 | 17.9 | 45.4
1,200 | 59.6 | 51.4 | 24.2 | 55.1
1,300 | 70.8 | 57.2 | 32.7 | 66.2
1,400 | 83.2 | 70.5 | 42.4 | 79.3
1,500 | 97.8 | 83.8 | 55.1 | 94.1
1,600 | 114 | 99.6 | 72.6 | 112
1,700 | 130 | 120 | 99.5 | 132
1,800 | - | 140 | 121 | 158
1,900 | - | 164 | 151 | 185
2,000 | - | 183 | 182 | 215
2,100 | - | 208 | 220 | 249
2,200 | - | 232 | 263 | 286
DISTRIBUTION A. Approved for public release; distribution unlimited.

## Slide 21: Laser Power
Sample Temperature [°C] | ER64-001 Laser Power Test 1 [W] | ER64-002 Laser Power Test 1 [W] | ER64-002 Laser Power Test 2 [W] | ER64-002 Laser Power Test 3 [W]
2,300 | - | 263 | 279 | 329
2,400 | - | 297 | 332 | 374
2,500 | - | 340 | 392 | 415
2,600 | - | 387 | - | -
DISTRIBUTION A. Approved for public release; distribution unlimited.

## Slide 22
DISTRIBUTION A. Approved for public release; distribution unlimited.

## Slide 23: ER64-002 Photo Comparison
DISTRIBUTION A. Approved for public release; distribution unlimited.
ER64002 pre-test
ER64002 post-test #1
ER64002 post-test #2
ER64002 post-test #3
ER64002 is visibly more reflective following HiTEMS tests

## Slide 24: Future HiTEMS Enhancements
Starting Summer 2023:
  Begin upgrade to a hemispherical emittance capability through remote operation/rotation of the sample mount in-vacuum
    Also provides ability to rotate to specific incident angles without re-pressurizing
    Near normal still available
Following hemispherical emittance upgrade:
  Combine the HiTEMS and Low Temperature Emissivity Measurement System (LoTEMS) systems
    Enables the full temperature range from ambient to 2,600°C without interrupting testing
DISTRIBUTION A. Approved for public release; distribution unlimited.

## Slide 25: Low Temperature Emissivity Measurement System (LoTEMS)
Not yet operational
Temperature range: ambient to approximately 1,300°C
Utilizes internal contact heating
Measures total and spectral near-normal emittance
Needs mirror upgrade to resemble HiTEMS
Needs Labview and hardware upgrade
DISTRIBUTION A. Approved for public release; distribution unlimited.