## Slide 1: UPWT Check Standard Model Vortex Characterization

## Slide 2: Background
Centerline vortex caused by presence of nozzle block T-rail discovered during CFD studies of UPWT in 2019
Asymmetric flap loading occurred during CobraMRV Control Surface Effectiveness test in 2020 while model was at 0° beta with symmetric flap deflection angles
  Asymmetric loading was believed to be caused by presence of vortex following additional CFD and flap check loads
Would an internal balance be able to detect the presence of the vortex as well?

## Slide 3: Objective
Devised an experiment using the UPWT Check Standard Model (CSM) to characterize centerline vortex using an internal balance
Move model to different locations within the test section to see changes in balance force and moment coefficients as model passes into and out of vortex
Compare changes in balance loads to force/moment coefficient uncertainty from previous CSM tests
Test usual CSM flow conditions along with flow conditions from CFD study

## Slide 4: Test Matrix
Sweep | Alpha | Beta | Roll | Model Nose
1 | sweep from -3° to 6° | 0° | 0° | centered in test section
*2 | sweep from -3° to 6° | +2°, +4° | 0° | centered in test section
3 | sweep from -3° to 6° | 0° | 0° | ±6 in.
4 | 0° | -3° to +3° | 90°, 270° | centered in test section
5 | 0° | -3° to +3° | 90°, 270° | off center
6 | 0° | 0° | 90°, 270° | -10 in. to +10 in.
*CFD comparison sweeps
Primarily focused on sweeps 1-3, 6
Notes: Sweeps 4-5 did not show much at time of testing, had to remove from matrix

## Slide 5: Support System Motion for Model Attitude

## Slide 6: Test Conditions
Condition | Mach | Ho (psf) | To (°F) | Re (1x106) | Q (psf) | P (psf) | Dewpt (°F)
1 | 2.60 | 1686 | 125 | 2.00 | 399.7 | 84.48 | -24.9
2 | 3.40 | 2737 | 150 | 2.00 | 335.0 | 41.39 | -16.4
3 | 4.30 | 4261 | 150 | 2.00 | 245.4 | 18.96 | -7.5
4 | 2.60 | 3371 | 125 | 4.00 | 799.5 | 169 | -12.2
5 | 3.40 | 5474 | 150 | 4.00 | 669.9 | 82.8 | -2.4
6 | 4.30 | 8521 | 150 | 4.00 | 490.7 | 37.9 | 6.2
7 | 2.399 | 2280.5 | 125 | 3.00 | 629.20 | 156.13 | -18.7
Conditions 1-6 are the usual CSM conditions
Condition 7 is one of the conditions from the previous CFD studies
Focusing on conditions 1–3 for presentation
Notes: Focusing on Conditions 1-3 for presentation

## Slide 7: Instrumentation
UT59A internal balance
7 x 5-psia Druck transducers
  4 base pressure measurements
  2 cavity pressure measurements
  1 tunnel static pressure measurement
Pitch mechanism Q-flex accelerometer
Component | Load | Sensitivity
Normal | 400 lbs. | 310.9537 (lb./mv/V)
Axial | 40 lbs. | 36.9301 (lb./mv/V)
Pitch | 1200 in-lbs. | 721.9365 (in-lb./mv/V)
Roll | 200 in-lbs. | 191.502 (in-lb./mv/V)
Yaw | 300 in-lbs. | 204.0489 (in-lb./mv/V)
Side | 100 lbs. | 102.2324 (lb./mv/V)

## Slide 8: Pitch Sweep Results

## Slide 9: AF Coefficient
Notes: Not a large difference in axial force trends, though Mach 4.3 shows a downward shift when off tunnel centerline

## Slide 10: NF Coefficient
Notes: NF tracks well

## Slide 11: SF Coefficient
Notes: SF coefficient at centerline increases more rapidly as Mach number increases, hard to tell if caused by vortex or not

## Slide 12: RM Coefficient
Notes: At lower Mach, RM coefficient is than when negative when on-centerline. More dispersion of data at lower Mach number. For higher Mach numbers, centerline data tracks off-center data well

## Slide 13: PM Coefficient
Notes: PM coefficient data appears to track well

## Slide 14: YM Coefficient
Notes: At lower Mach number, YM coefficient data appears to be more spread apart, but tightens up as Mach number increases

## Slide 15: Lateral Sweep Results

## Slide 16: AF Coefficient
Notes: AF coefficient data tracks well

## Slide 17: NF Coefficient
Notes: Trends in NF coefficient appear to be mirrored as model moves laterally across test section, with a bump in the center at Mach 2.6. Weird things happening at sides around +10 in. Maybe vertical tail is catching sidewall vortex? NF trends are closer together at Mach 2.4, bubble at centerline appears to flatten as Mach number increases

## Slide 18: Vortex Overlay - NF
Notes: Trends in NF coefficient appear to be mirrored as model moves laterally across test section, with a bump in the center at Mach 2.6. Weird things happening at sides around +10 in. Maybe vertical tail is catching sidewall vortex? NF trends are closer together at Mach 2.4, bubble at centerline appears to flatten as Mach number increases

## Slide 19: SF Coefficient
Notes: Bubble roughly at centerline forms in data at Mach 2.4, but dissipates as Mach number increases

## Slide 20: Vortex Overlay - SF
Notes: Bubble roughly at centerline forms in data at Mach 2.4, but dissipates as Mach number increases

## Slide 21: RM Coefficient
Notes: RM coefficient trends track each other well. Data crosses at Mach 2.4 near centerline and not at other Mach numbers. Hard to tell if caused by presence of vortex

## Slide 22: PM Coefficient
Notes: PM coefficient tracks well at all Mach numbers. Some symmetry present as model moves across test section. Not seeing much influence from vortex

## Slide 23: YM Coefficient
Notes: Large bubble in data at Mach 2.4, right near centerline. Bubble dissipates as Mach number increases

## Slide 24: Vortex Overlay – YM
Notes: Large bubble in data at Mach 2.4, right near centerline. Bubble dissipates as Mach number increases

## Slide 25: Uncertainty: Off-Centerline Pitch Sweep Comparisons

## Slide 26: Mach 2.6 Re∞ 2.0 Million/ft
3 repeated pitch sweeps at tunnel centerline
0, +6, -6 in. pitch sweeps
α (deg)
α (deg)
α (deg)
α (deg)
α (deg)
α (deg)
α (deg)
α (deg)
α (deg)
α (deg)
α (deg)
α (deg)
Notes: Uncertainty increases for normal force and pitching moment coefficients as model moves off centerline, but data is still within bounds

## Slide 27: Mach 2.6 Re∞ 2.0 Million/ft
3 repeated pitch sweeps at tunnel centerline
0, +6, -6 in. pitch sweeps
α (deg)
α (deg)
α (deg)
α (deg)
α (deg)
α (deg)
α (deg)
α (deg)
α (deg)
α (deg)
α (deg)
α (deg)
Notes: Side force, rolling moment, and yawing moment coefficient uncertainty all increase when off centerline

## Slide 28: Mach 3.4 Re∞ 2.0 Million/ft
3 repeated pitch sweeps at tunnel centerline
0, +6, -6 in. pitch sweeps
α (deg)
α (deg)
α (deg)
α (deg)
α (deg)
α (deg)
α (deg)
α (deg)
α (deg)
α (deg)
α (deg)
α (deg)
Notes: Increase in uncertainty appears to be less compared to Mach 2.6 cases

## Slide 29: Mach 3.4 Re∞ 2.0 Million/ft
3 repeated pitch sweeps at tunnel centerline
0, +6, -6 in. pitch sweeps
α (deg)
α (deg)
α (deg)
α (deg)
α (deg)
α (deg)
α (deg)
α (deg)
α (deg)
α (deg)
α (deg)
α (deg)
Notes: Larger increase for side force, rolling moment, and yawing moment coefficients than for other coefficients

## Slide 30: Results and Discussion
AF Coefficient
  Pitch sweeps
    No large differences at each Mach number, though slight drop in axial force when off-centerline at Mach 4.3
  Lateral sweeps
    No large change between sweeps
NF Coefficient
  Pitch sweeps
    No large change between sweeps
  Lateral sweeps
    Trends appear to be mirrored as model moves laterally across test section, with a “bump” occurring near centerline at Mach 2.6
    Bump likely caused by presence of vortex at centerline of test section
    Bump near tunnel centerline flattens as the Mach number increases
    As Mach number increases, the nozzle block moves forward exposing less of the T-rail. This causes the vortex pair to diminish in size and strength
SF Coefficient
  Pitch sweeps
    SF increases more rapidly at centerline as Mach number increases
    Possible indication of vortex effecting increase in SF, but tricky to identify cause
  Lateral sweeps
    Bump occurs at centerline for SF data as well which flattens as Mach number increases, indicating reduction in size and strength of vortex

## Slide 31: Results and Discussion
RM Coefficient
  Pitch sweeps
    More dispersion of data at lower Mach numbers
    For higher Mach numbers, centerline data tracks off-center data well
  Lateral sweeps
    Likely no influence of vortex
PM Coefficient
  Pitch sweeps
    No large change between sweeps
  Lateral sweeps
    No large change between sweeps
YM Coefficient
  Pitch sweeps
    Data appears to be more spread apart, but tightens up as Mach number increases
  Lateral sweeps
    Bump occurs at centerline for SF data as well which flattens as Mach number increases, indicating reduction in size and strength of vortex
Uncertainty
  Higher uncertainty in data as model moves off centerline

## Slide 32: Future Work
Analyze beta (CFD) sweeps
Continue uncertainty analysis for pitch sweeps at Mach 4.3
Analyze uncertainty for lateral sweeps at all Mach numbers
Plan to add similar sweeps to next check standard test for additional data supporting facility modifications

## Slide 33: Acknowledgements
Bryan Falman
Jim Ross
Jose Mondragon

## Slide 34
Questions?