## Slide 1
PACE OCI crosstalk characterization based on pre-launch testing
Shihyan Lee1,2, Gerhard Meister1, Samuel Kitchen-McKinley3, Joseph Knuble4, Ulrik Gliese5, and Robert Bousquet6
1NASA Code 616, Ocean Ecology Lab, Goddard Space Flight Center, Greenbelt, Greenbelt Rd., Maryland, MD 20771, USA
2Science Applications International Corporation, 12010 Sunset Hills Rd., Reston, VA 20190, USA
3Science Systems and Applications Inc., 10210 Greenbelt Road, Lanham, Maryland 20706, USA
4NASA Goddard Space Flight Center, Greenbelt, Greenbelt Rd., Maryland, MD 20771, USA
5KBR, 7701 Greenbelt Rd., Greenbelt, MD 20770, USA
6Genesis Engineering Solutions, Inc., 4501 Boston Way, Lanham, MD 20706, USA
SPIE Sensors + Imaging, Amsterdam, NL September 6, 2023
Notes: 1
1

## Slide 2
Source: collimated, mono-chromatic light (GLAMR) @ 583 nm.
Near-field crosstalk into adjacent pixels and wavelengths (bands); exponential decay
Ghost at specific wavelengths and pixels; constant
High Contrast Scene Measurements
OCI
GAToR
LSF Bench
1 SP mask
Straylight
Curtain
GLAMR
ghost@ 382.8nm
31 pixel wide
Wavelength (nm)
pixel
Near-field crosstalk
LSF = Line spread function

## Slide 3
Ghost and TDI are spatially asynchronous. Source and corresponding ghost pixels are at mirrored positions.
Effects illustrated by OCI optical model
Quantified using test data
Ghost vs. TDI
(a)
(b)
(c)
(d)
ghost
@pos=16
ghost
@pos=10
ghost
@pos=1
ghost
@pos=5
Wave (nm)
pixel
(a)
(b)
(c)
(d)

## Slide 4
LSF 1 science pixel mask (LSF-SP) vs. 7 science pixel mask (LSF-7SP)
Stronger and wider ghost in LSF-7SP data than LSF-SP data
High Contrast Scene Measurements
LSF-7SP
ghost
@382.8nm
LSF-SP
ghost
@382.8nm
(a)
(b)
Wavelength (nm)
pixel
Wavelength (nm)
pixel
PSF = point spread function, g_wave = GLAMR wavelength

## Slide 5
Ghost Magnitude: LSF-7SP ~ 7x LSF-SP
Ghost Spatial extent: LSF-7SP ~ LSF-SP + 6 pixels
Ghost = convolution of (source and LSF-SP ghost feature)
OOB response (RSR): LSF-7SP ~ LSF-SP
Ghost: LSF-SP vs LSF-7SP
LSF-7SP
LSF-SP
OOB
OOB
(b)
(a)
ghost 0.22%
ghost 0.031%
37 pixel
31 pixel
OOB = out-of-band

## Slide 6
From LSF-SP: Ghost = convolution of (source and LSF-SP ghost feature)
In-band response: source profile
Ghost response: gaussian like shape
OOB (not shown) response = in-band response
Ghost: monochromatic broad source
CAL1: GLAMR
@575nm
Wavelength (nm)
pixel
In-band
Ghost
@400nm
dn

## Slide 7
LSF: limited wavelengths tested; Most accurate in ghost characterization
GLAMR: data in 1-nm step; LFS data Gap filling
Ghost characterization
(a)
(b)
(nm)
(nm)
xc wave = ghost wavelength, xc val = ghost magnitude

## Slide 8
Crosstalk Map Method
PSF
  Determine ghost wavelength and magnitude in LSF and CAL data
  Interpolate from measured wavelength to all CCD bands
  In LSF-SP data, remove ghost in PSF to create nominal PSF.
  For each band, compute modeled nominal PSF by interpolating 2 adjacent nominal PSFs.
  Add ghost based on the predicted wavelength and magnitude.
PSF to crosstalk
  For each band, compute crosstalk map with sender bands from the same CCD using the modeled PSFs.
  Crosstalk map is computed in 0.625nm resolution from 325nm to 890 nm; aggregated to 5nm bands for science data analysis

## Slide 9
Model PSF
Nominal PSF (Top R): ghost feature removed from measured data.
Modeled nominal PSF (Bot L): interpolated from 535 and 555 nm nominal PSF.
Modeled PSF (Bot R): Add ghost feature in modeled nominal PSF.

## Slide 10
Crosstalk
Compute crosstalk from modeled PSF
2 ghosts in crosstalk vs. 1 ghost in PSF.
2 bands send ghosts into the same wavelength
Nominal PSF

## Slide 11
PSF vs. crosstalk map

## Slide 12
Crosstalk map evaluation
WL-model = modeled white light PSF
WL-measure = measured white light PSF
GLAMR = monochromatic light PSF.
Estimate white light response from modeled PSFs.

## Slide 13
Summary
The crosstalk of the OCI blue CCD bands has a ghost feature.
The crosstalk including ghost feature were characterized using high-contrast scene (LSF), and broad source calibration (CAL-GLAMR) data.
The ghost has a spatially continuous impact up to +- 15 pixels from the source; impact can not be mitigated by typical cloud mask
On-going:
New ghost feature discovered in red CCD bands (impact a few bands ~880 nm from sources ~910 nm); characterization underway.
implement crosstalk correction in L1B processing:
  Crosstalk maps -> crosstalk correction LUT;

## Slide 14
backup

## Slide 15
ghost wavelength (Left) and magnitude (Right)
Black: model, red: FU1 data
FU1 measured vs. model

## Slide 16
Smaller width@95% max, used to identify ghost
Ghost signal is the convolution of source and TDI pixels

## Slide 17
LSF-7SP @ 360nm (cds8x1)
OOB + ghost
OOB + ghost
OOB

## Slide 18
Higher straylight/crosstalk at lower wavelengths
LSF-SP GLAMR
wave
pixel

## Slide 19
Higher straylight/crosstalk at lower wavelengths
Interpolate CCD straylight from GLAMR measured wavelengths.
475 nm measured vs. interpolated (linear)
LSF-SP GLAMR
Area of
Interest
Area of
Interest