## Slide 1
Jeremy Werdell, PACE Project Scientist
Robby Estep, OCI Project Manager
Gerhard Meister, OCI Instrument Scientist
Sentinel-3 Next Generation
Optical Mission Advisory Group Meeting No. 2
PACE OCI Lessons Learned
(through Phase C & into Phase D)

## Slide 2
NASA Plankton, Aerosol, Cloud, ocean Ecosystem (PACE) mission
https://pace.gsfc.nasa.gov
@NASAOcean
PACE will extend key systematic ocean color, aerosol, & cloud climate data records.
PACE will reveal the diversity of organisms fueling marine food webs & how ecosystems respond to change.
Looking at the ocean, clouds, and aerosols together will improve knowledge of the roles each plays in our planet.
Mission update:
Observatory full assembled & integrated
Mechanical (acoustics, vibration) testing complete
Thermal vacuum testing complete
Day(s)-in-the-life, etc. data flow testing in progress
Delivery to Kennedy Space Center in November 2023
60-day in-orbit commissioning to follow launch
The PACE Ocean Color Instrument (OCI):
340-890 nm @ 5 nm resolution in 2.5 nm spectral steps
Plus 940, 1038, 1250, 1378, 1615, 2130, & 2250 nm
2-day global coverage; 1-km2 @ nadir; ±20o fore/aft tilt
Performance driven by ocean color science requirements
2 contributed multi-angle polarimeters:
HARP-2 (UMBC)
  4 visible-NIR bands
  Wide swath; 2.5 km @ nadir
  Hyper-angular
  Cloud capabilities beyond OCI
SPEXone (SRON/Airbus)
    Hyperspectral UV-NIR
    Narrow swath; 3 km @ nadir
    5 angles
    Aerosol capabilities beyond OCI
Key characteristics:
9 January 2024 launch
676.5 km altitude
Polar, ascending, 98o
Sun synchronous
13:00 Equatorial crossing

## Slide 3
ocean color & the ocean color instrument
ocean color retrievals drive OCI’s design & performance requirements
hyperspectral scanning radiometer
(320) 340 – 890 nm, 5 nm resolution, 2.5 nm steps+
plus, 940, 1038, 1250, 1378, 1615, 2130, and 2250 nm
single science pixel to mitigate image striping
1 – 2 day global coverage (±56° cross track)
ground pixel size of 1 km2 at nadir
± 20o fore/aft tilt to avoid Sun glint
twice monthly lunar calibration
daily on-board solar calibration
<0.5% total system error for VIS-NIR
SNRs optimized for ocean color science
average data rate of 20 Mbps
mass of ~260 kg CBE (includes portion of tilt structure)
simulated top-of-atmosphere data available
+ with 1.25 nm steps in several spectral regions
* developed primarily for mechanical processing assessments

## Slide 4
general outline
Project (Mission) Science met weekly or more frequently with Systems Engineering
HQ directed baseline requirements ownership to the Project
Project developed simulators & new tools for requirements verification
OCI’s highly integrated (interdependent design) created challenges
Project practiced their (science & calibration) data analysis tools on the ETU
Instrument team developed a high-fidelity “virtual instrument” model

## Slide 5
caveat

## Slide 6
a science-engineering Rosetta Stone
Project Science and Systems Engineering started weekly+ standing meetings just after mission designation
Examples of successes realized by the trust & familiarity gained through these interactions:
Facilitated flexibility to adjust requirements as necessary*
  reduced the dynamic range requirement for chlorophyll-a bands to obtain higher SNRs
  reduced throughput requirements for the blue focal plane assembly when choosing attenuators
Enabled natural collaboration to pursue and resolve technical problems
  To address the SWIR bands’ extended impulses:
    Science realized an algorithmic correction
    Systems Engineering developed a new mechanism
    Both iterated on test requirements to support the above
Kept focus on design driving requirements (& identifying what “science really needed”)
* more on this next

## Slide 7
the Project owned baseline requirements
HQ imposed a “design-to-cost” (DTC) paradigm, which ultimately allowed the Project to own baselines
(In the end) DTC provided design flexibility & requirements flexibility simultaneously
  Project designed to HQ-held Level-1 threshold requirements with margin
  Project managed Level-1 baseline and Level-2 requirements
  allowed Project to manage design evolution & hardware implementation

## Slide 8
requirements at mission confirmation
Ground sample distance of 1 ± 0.1 km2 at nadir
Equatorial crossing between 11:00 to 13:00 local time
1-day to solar zenith ≤ 75o & sensor zenith ≤ 60o
Sun glint mitigation (<8% of the 2-day global coverage missed)
Twice-monthly characterization of instrument stability, including lunar observations through the earth viewing port that illuminate all detector elements
Daily characterizations of OCI instrument changes using an independent, on-board capability that illuminates all detector elements
An OCI with center wavelengths from 350-865 nm at 5 ± 1 nm resolution
An OCI with bands centered on 940, 1038, 1250, 1378, 1615, 2130, and 2260 nm at atmospheres and ocean performance levels
Multi-angle polarimeter(s)
36-month mission life
Ground sample distance of 1 ± 0.1 km2 at nadir
Equatorial crossing between 11:00 to 13:00 local time
2-day to solar zenith ≤ 75o & sensor zenith ≤ 60o
Sun glint mitigation (<9% of the 2-day global coverage missed)
Once-monthly characterization of instrument stability, including lunar observations through the earth viewing port that illuminate all detector elements
Daily characterizations of OCI instrument changes using an independent, on-board capability that illuminates all detector elements
An OCI with center wavelengths from 350-820 nm at 5 ± 1 nm resolution
An OCI with bands centered on 865, 940, 1250, 1378, 1615, 2130, and 2260 nm at atmospheres performance levels
Multi-angle polarimeter(s)
18-month mission life
Data Product | Uncertainty
Water-leaving reflectances from 350 to 400 nm | 0.0083 or 30%
Water-leaving reflectances from 400 to 600 nm | 0.0024 or 6%
Water-leaving reflectances from 660 to 710 nm | 0.00084 or 12%
Total aerosol optical depth at 380 nm | 0.06 or 40%
Total aerosol optical depth at 440, 500, 550 and 675 nm over land | 0.06 or 20%
Total aerosol optical depth at 440, 500, 550 and 675 nm over oceans | 0.04 or 15%
Fine mode fraction of aerosol optical depth over oceans at 550 nm | ±25%
Cloud layer detection for optical depth > 0.3 | 40%
Cloud top pressure of opaque (optical depth > 3) clouds | 60 hPa
Optical thickness of liquid clouds | 25%
Optical thickness of ice clouds | 35%
Effective radius of liquid clouds | 25%
Effective radius of ice clouds | 35%
Data Product | Uncertainty
Water-leaving reflectances from 350 to 400 nm | 0.0057 or 20%
Water-leaving reflectances from 400 to 600 nm | 0.0020 or 5%
Water-leaving reflectances from 660 to 710 nm | 0.0007 or 10%
Total aerosol optical depth at 380 nm | 0.06 or 40%
Total aerosol optical depth at 440, 500, 550 and 675 nm over land | 0.06 or 20%
Total aerosol optical depth at 440, 500, 550 and 675 nm over oceans | 0.04 or 15%
Fine mode fraction of aerosol optical depth over oceans at 550 nm | ±25%
Cloud layer detection for optical depth > 0.3 | 40%
Cloud top pressure of opaque (optical depth > 3) clouds | 60 hPa
Optical thickness of liquid clouds | 25%
Optical thickness of ice clouds | 35%
Effective radius of liquid clouds | 25%
Effective radius of ice clouds | 35%
threshold L1 requirements
baseline L1 requirements

## Slide 9
hard conversations about descopes
what did DTC look like on the OCI design at mission confirmation (KDP-C)?
( what are the baseline capability descopes the mission could have taken to maintain cost/schedule? )
UVVIS & VISNIR Focal Plane Assemblies (FPA) (not pictured)
Relax SNR requirements
Relax absolute calibration requirements
Relax relative calibration requirements
SWIR Detection Assembly
Descope from 32 FPAs to 12 FPAs
Descope from 16 fibers to 6 fibers
Relax SNR requirements
Relax absolute calibration requirements
Relax relative calibration requirements
Solar Calibration Assembly
Descope from 3 calibration targets to 1 calibration target
Remove carousel mechanism
Optical System
Remove bands between 320nm – 340nm
Tilt System (not pictured)
Increase allowable time to tilt and settle

## Slide 10
new simulators & tools for requirements verification
Project Science invested in development of (coarse) OCI Level-1B simulator, as well as an uncertainties propagation model, which in combination enabled rapid requirements verifications
General flow (details on next slide)
  identify target top level science requirements (e.g., Level-1 water leaving reflectances uncertainties)
  define top level instrument-level requirements (e.g., Level-2 uncertainty of TOA radiances)
  derive uncertainty of TOA radiances from (e.g., Level-3) pre-launch characterization measurements*
  combine the measured/estimated uncertainties into the instrument-level uncertainty
  evaluate if top level science requirements have been met
Tools
  OCI simulator: pyTOAST ( https://oceancolor.gsfc.nasa.gov/data/pace/test-data/ )
  uncertainties propagation model: Zhang et al. 2022; doi:10.1364/OE.460735
* we did an imperfect job, but aimed for reasonable estimates & assumptions (error estimation is always tricky, so no possibility for perfection)

## Slide 11: Requirement Connection & Verification Example
Level-1 science data product (L1RD-##)
Water-leaving reflectance's* with uncertainty of 0.002 (unitless)
* wavelengths not listed for brevity
Level-2 (MRD-##*)
Water-leaving reflectance uncertainty allocations (sum in quadrature to ≤ 0.002 (unitless)):
* L1RD-## flows to 32 MRD requirements
Systematic error | Random error | Model error
0.0018 | 0.0006 | 0.0005
Level-3 (OCI-##*)
OCI shall meet SNRs with a uniform spectral radiance in all science pixels (flat field) at Ltyp at the bandwidth as defined [in Table 2] for all bands
* MRD-## flows to 25 OCI requirements
Instrument performance data / models
OCI total random & systematic error CBEs
Total error allocations
Derived L1 uncertainties
Combine allocations
Combine with model error
Apply to simulated OCI TOA radiances
This “artifacts-to-imagery” simulator allowed rapid assessment of OCI performance & aids in identifying margin in Level-3 and above
engineering
Our mission Standing Review Board praised this capability at both our Critical Design Review & Systems Integration Review
Notes: Uncertainties defined for ~15 nm aggregated bands (not for every 5 nm)
Model error fixed on heritage algorithm – provides sandbox in which OCI can play

## Slide 12
addressing a highly integrated system
12
SCA
MOB
MOSB
BLUE FPA
RED FPA
MLA/Fiber
UVNIR LHPs/ Radiators
SDA (Procurement From Space Dynamics Lab)
ICDU
MCE
DAU
Structure
RSS, ES & Support Structure
OCI’s design is highly integrated & interdependent
The issues:
  Interfaces are intertwined, not decomposed or easily verified until system level
  All components have complex operations
  Adds increased risk to higher levels of assembly
OCI
FSW
What is a highly integrated system?

## Slide 13: The OCI testing spiral – decomposition, test planning challenges, & increased risk (1/2)
I can’t verify the FPA performance without the DAU… and harness length meters
May as well add the structure to hold it together
The MOB rotating telescope actuator and half angle mirror actuator requirements need optics for verification
I need an MCE Ebox to drive the actuators
Measuring optical performance while the telescope is spinning requires an FPA that is TDI sync’d
The FPA doesn’t TDI sync without the DAU
The MOB and MOSB optical and mechanical design are tightly coupled - the alignment of one impacts the other
If I have an MCE and DAU, I may as well add the ICDU and interface through I&T/ITOS
High Risk
Due To Late Verification &
Validation
The optical system doesn’t have an obvious breakpoint for verification
Need independent thermal control zones for calibration
Example of the circular conversation the engineering team navigated writing L4 requirements, decomposing, and test planning

## Slide 14
navigating test-as-you-fly challenges for subsystem development
Optical Module (OM)
UVNIR Focal Plane Assembly (FPA)
Data Acquisition Unit (DAU)
SWIR Detection Assembly (SDA)
“Ideal” Test Plan: Rotating image at all wavelengths
Challenges: Fast image speed and GSE readout speed
Actual Test Plan: Static optical image measurements w/limited wavelength/bands
Risk: Image blur, MTF, scatter/ghosts/SL
“Ideal” Test Plan: DAU power w/moving image
Challenges: Out vs In-House Subsystems
Actual Test Plan: GSE power and static image
Risk: EMI/EMC, Telescope and TDI synchronization, SIDECAR data processing
“Ideal” Test Plan: DAU backend and moving image
Challenges: Organizational boundaries, optical GSE complexity and cost
Actual Test Plan: GSE backend & static optical image
Risk: EMI/EMC, timing, data processing
“Ideal” Test Plan: FPA frontend
Challenges: Organizational boundaries
Actual Test Plan: ETU FPA frontend
Risk: EMI/EMC, timing, data processing
Key Lesson Learned: Manage the camera (DAU-FPA) as one system with one PDL

## Slide 15: The OCI testing spiral – decomposition, test planning challenges, & increased risk (2/2)
We didn’t know if the OCI architecture/design would work until 1 month before mCDR when we got end to end 1st light
Photos contained in this album have been reviewed by the Export Control Office and are released for public view

## Slide 16
lessons learned from OCI’s highly integrated system
Advocate for a program with ETU hardware & end-to-end testing* as soon as possible
  enabled real results & discussions leading up to & at our Critical Design Review & beyond
  required being flexible with ETU hardware, which was limited in scope
  required general flexibility & patience, as well as the ability to iterate routinely
Maintain parallel paths on hardware procurements (when practical)
  gave the Project a jump on many critical aspects of OCI development & testing
    GSE, facilities, processes, personnel requirements
    vendor selection & contract implementation
  demonstrated incorrect assumptions
    materials, facilities, or required expertise not available
    expected capabilities lost or not present
    what we thought we needed isn’t what we actually needed
  identified risks that demonstrated need(s) for alternative paths/providers
  some benefits did not require flight-like hardware
Build a strong modeling team* with independent checks & reviews
* more on these later

## Slide 17
ETU data analysis tools
Developing & exercising (practicing using our) data analysis tools on the ETU proved critical
Budgeting for a comprehensive characterization campaign also proved critical
Flight I&T included little time for analysis
  time crunch during flight testing made easier knowing what to expect from ETU testing
  allowed for improving GSE issues, even when data not yet fully evaluated
  enforced early development of scientific analysis tools for rapid test interpretation (go/no-go)
  would not have otherwise known how to understand anomalous results
    still, ETU testing arriving late prohibited time to adjust the SWIR electronics to improve hysteresis & band-to-band registration
Recognize early the complex interrelation between performance verification, calibration, and GSE
  separate GSE & instrument sources of uncertainty
  understand absolute vs. relative requirements, correlations, & random vs. systematic contributors
  identify what can be measured pre-launch & how on-orbit capabilities influence the test plan
  as an aside: most NASA instrument evolution is defined by uncertainty improvements
    13 of 53 OCI Level-3 performance requirements are residual uncertainty requirements after a calibration correction is applied
For a systems engineer working in calibration, a working knowledge of “uncertainty” in terms of scientific measurements is important.
  Most NASA instrument evolution is defined by uncertainty improvements.
  13 out of 53 OCI L3 performance requirements are residual uncertainty requirements after a calibration correction is applied.
  Which error sources are correlated and which aren’t?
    A.k.a can I just RSS these terms?

## Slide 18: pictures of GSE used for OCI calibration
20’’ White Light Sphere & Polarizer
Limited Performance Test GSE
Line Spread Function (LSF) and Stray Light
GLAMR Tunable Laser
GAToR Mechanical Positioning System
GAToR PDR
Nikon Image of Slanted Reticle
OCI Image of Slanted Reticle
GAToR CDR
Photos contained in this album have been reviewed by the Export Control Office and are released for public view

## Slide 19: pictures of OCI real time analyses tools developed for OCI

## Slide 20
a virtual instrument model
Investment in a “virtual” OCI instrument model (Matlab) helped tremendously in making assessments & architectural decisions as the Project matured & validated the instrument
Development of this virtual instrument model supported OCI’s design
  OCI’s band-to-band registration issue was simulated, then used in an impact analysis
  FPA results were simulated that otherwise would have remained puzzling
In principle, this model will support performance diagnostics after launch

## Slide 22
general backup

## Slide 23
ORCA Proposal
PACE/ORCA TMR
OCI Pre-Phase A
OCI TMC
PACE MCR
OCI SRR
OCI DTC Design Evolution
PACE Start

## Slide 24
OCI PDR
OCI CDR
Tilt Trades – Post SRR
OCI DTC Design Evolution

## Slide 25: OCI Instrument Overview - Backup
Instrument Mass | 260 kg CBE , 305 kg NTE
Instrument Power | 280 W CBE , Orbital Average
Thermal System | Passive Cooling with Loop Heat Pipes
Mechanisms | Rotating Telescope Mechanism (~6Hz) Half Angle Mirror Mechanism (~3Hz) Solar Calibration Mechanism (5000 Cycle)
Deployments | Rotating Telescope Launch Lock
Ground Sample Distance | 1050m at Nadir
Aperture | 90mm
iFOV | 0.08o Along Track x 1.42o Cross Track
UVNIR System EFL | 0.134m
UVNIR System f/no | 1.49
Total Field of Regard | +/- 56.5o
Dynamic Range | SNR at Ltyp , No Saturation at Lmax
UVNIR Bands (nm) | 342.5nm – 887.5nm , 5nm Resolution
NIRSWIR Bands (nm) | 940, 1038, 1250, 1378, 1615, 2130, 2260
UVVIS & VISNIR FPAs | 2 CCDs, 128 x 512, 26 micron Integrated 14 Bit ADC
NIRSWIR FPAs | 16MCT/16InGaAs Photodiodes, < 250 micron analog output to SIDECAR ASIC
Relative Radiometric Accuracy | < 0.5% Pre-Launch 1-sigma
SNR @ Ltyp | > 1000:1: 340 – 700nm > 600:1: 700 – 865nm > 50:1 940 – 2260nm
On-Board Solar Calibration Assembly | Daily & Monthly Solar Calibration Targets
Orbital Average Data Rate | 13 Mbps up to 40 Mbps
Cross-Track Rotating Telescope
UVVIS & VISNIR Slit Grating Hyperspectral Spectrographs
  2 CCD FPAs with On-Chip TDI Synchronized with Rotating Telescope and Half Angle Mirror
NIR-SWIR Fiber Coupled Multiband Filter Spectrograph
  32 InGaAs/MCT Single Photo Diode FPAs with Digital TDI
Concept follows the heritage of the SeaWiFS, MODIS, and VIIRS
+Z
+X
+Y

## Slide 26: What is OCI? - Subsystem Description (1/2)
+Y
+Z
+X
Optical Module
Main Optical Bench (MOB)
  Rotating Telescope Assembly
  Half Angle Mirror Assembly
Main Optical Sub-Bench (MOSB)
  Slit Grating UVNIR Hyperspectral Spectrograph
  Slit - Fiber Coupled NIR SWIR “Light Bucket”
Solar Calibration Assembly (SCA)
Integrated carousel mechanism with 3 calibration targets and reusable door
SWIR Detection Assembly (SDA)
NIR-SWIR (0.94um – 2.2um) Fiber Coupled Multiband Filter Spectrograph
32 InGaAs/HgCdTe Photo Diode FPAs with Digital TDI
Integrated SIDECAR ASIC (SSM)
Instrument Deck Structure
Electrical Interface

## Slide 27: What is OCI? - Subsystem Description (2/2)
+Y
UVNIR (0.34um – 0.87um) Focal Plane Assemblies (FPAs)
CCD with On-Chip TDI and Front End Electronics w/Integrated ADC
Synchronized with Rotating Telescope and Half Angle Mirror
+Z
Data Acquisition Unit (DAU)
CCD Clock and Bias
SDA Power
Data Processing
Motor Control Electronics (MCE)
RT, HAM, SCA Actuator Control
Instrument Command & Data Unit (ICDU)
Processor, House Keeping, Thermal Control, Power Distribution
Radiator Earth Shield
UVNIR FPA Loop Heat Pipes

## Slide 28: Leverage parallel paths to mitigate and spread out the over-constraint risk (1/3)
Parallel paths are a good thing!
  Used to mitigate technology & performance risk (Early Project Phases A-B)
    Can help lock in the design architecture and evaluate different technologies
    Enables diverse engineering to determine the right architecture and reduce unknown risks in the later project phases
  Used to mitigate procurement schedule delay & quality control risks (Later Project Phases C-D)
    Multiple vendors building same h/w to mitigate risk – keeps the project from stopping due to single vendor delays - keeps options open
  OCI parallel paths examples to mitigate proposal inadequacies (pre-phase A to phase B)
    Optical System Design – 3 options
    NIRSWIR Photodiodes – Many Options
  OCI parallel path examples to reduce schedule & quality control risks
    Examples: LHPs (ETU Upscreens), Lens assembly (ruda vs corning), mechanisms (honey bee vs thin gap)
    Even if simple, get h/w in quickly!
The downsides of parallel paths
  More to designs and potential interfaces to manage
  Costs more money upfront & requires more management and tracking
  Can confuse the team who might not see the whole picture - communication is key
Parallel path decision making is not straight forward – See the next few charts on OCI lessons and how to deal with these questions
  When do I close a trade to make progress?
  When do I keep design space open?
  How do I mange a team with many design options through parallel paths?

## Slide 29: Parallel Path Example – OCI Optical Architecture & Design (2/3)
OCI solution space at PDR had just down selected from 4 options to 1
The lesson is to stay patient and open minded – set up the project in the right design space
  Even after ORCA proposal and hardware development – the design still wasn’t where it needed to be
  It was very tempting to “plow ahead” to make “progress”
Configuration 1: Full in-house optics design (front-end and back-end)
  Minimum disruption to current design
  Red Channel: not viable
  Blue Channel: looks viable, with possible degraded performance
Configuration 1A: Full in-house with biconic surfaces
  Minimum disruption to opto-mechanical design
  Additional fabrication and alignment challenges
  Red Channel: no solution found
  Blue Channel solution found
Configuration 2: In-house front-end + vendor back-end
  Slight disruption to opto-mechanical design
  Red Channel: solution found
  Blue Channel: solution found
Configuration 3: TNO front-end and back-end
  Major disruption to OCI architecture and opto-mechanical design
  Red Channel: solution found
  Blue Channel: solution found

## Slide 30: Parallel Path Example – InGaAs vs HgCdTe Detectors for 0.94um – 2.2um (3/3)
OCI solution space at CDR had just down selected from 6 options to 2
A few months before CDR PACE OCI was at risk of violating L1 threshold requirements
  HgCdTe diodes tested to date had a input frequency dependent non-linearity that produced uncertainties greater than requirement
Photodiode detector lessons learned
  Maintain consistent detector package interface – to enable late swapping
  Keep lots of EEE parts (JFET in this case) on hand
  Have 2 independent groups verify the measurement/performance (Vendor/SDL/GSFC)
  TAYF – Don’t overfill a photodiode, keep the illumination spot flight like.
    There are many technical details to share – just ask offline
  Don’t stop talking to vendors and looking for options until you have the flight solution
  Stay patient & communicate the risk
Pre-CDR detector trade space (photos omitted to stay ITAR free)
  Discovery InGaAs
  TJT LPE (Liquid Phase Epitaxy) HgCdTe
  TJT MBE (Molecular Beam Epitaxy) HgCdTe
  Discovery Extended InGaAs
  TJT InGaAs: 0.94um – 1.38um
  LMW MOCVD (metal organic chemical vapor deposition) HgCdTe: 1.6um – 2.2um

## Slide 31: Design Evolution of OCI SWIR - Good Architecture Decisions
Around SRR time frame, ~Feb 2017, OCI decided to break off the SWIR bands with fiber optics
Free Space Optics SWIR Architecture
Highly Integrated
Alignment “Nightmare” with little lower level verification
More Cable and Harness on the Optical Module
Fiber Optics SWIR Architecture
Cleaner Optical Interface
Enable easier lower level verification
Outsourceable, reduce in house burden
Became SWIR Detection Assembly (SDA) built by the Space Dynamics Lab in Logan Utah
SWIR Design Evolution Allowed for simpler OCI Interface
Fiber Interface

## Slide 32: OCI Decision to Move the SDA Out-Of-House
In-house design enabled OCI to evaluate whether to develop the SDA in-house or to procure
Initially the SDA was planned to be an in-house development effort
  A full in-house design was developed
    Helped OCI understand the level of effort/complexity and that it would be difficult for the current team to build both the UVNIR and the SWIR Instrumentation while meeting schedule
    The design evolution (ex. fiber interface) allowed for a clean separation of SWIR functions and hardware
    Generated and Request For Information (RFI) to understand whether vendor capabilities existed and to get the associated cost and schedule estimates
    Generated Request Proposal (RFP) once the decision was made to go out-of-house and to compare vendors
    The in-house design effort was extremely beneficial for evaluating RFI/RFP responses and for generating the associated Spec and SOW documents