In-Situ XRD/XRF to Support Life Detection on Mars

E. B. Rampe1, D. F. Blake2, P. Sarrazin3, T. F. Bristow2, M. Gailhanou4, B. Lafuente5, V. M. Tu6, K. Zacny7, R. T. Downs8, 1NASA JSC, 2NASA Ames, 3eXaminart LLC, 4CNRS – Universite Paul Cezanne, 5SETI, 6Jacobs Jets-II NASA JSC, 7Honeybee Robotics, 8Univ. Arizona.

X-ray diffraction / X-ray fluorescence (XRD/XRF) analysis provides the most comprehensive mineralogical / compositional characterization of rocks and soils of any flight-capable technique. XRD data provide quantitative mineralogy (including abundance of X-ray amorphous materials) and crystal chemistry (structure, elemental composition and valence state), and XRF data provide complimentary major, minor, and some trace element abundances. Both types of data are important in evaluating habitability (environment of formation) and biosignature preservation/degradation (post-depositional diagenetic change). Whether or not a relict biosignature is detected, the mineral assemblage and its geochemistry can be used to determine the habitability of an ancient environment (e.g., salinity, pH, temperature), and to identify potential sources of energy for life (e.g., elements in different redox states). In this respect, a null result (a habitable environment lacking evidence of life) can play an important role in constraining the parameters of the search. Conversely, diagenetic alteration (taphonomic change) resulting from post-depositional variations in temperature, pressure or fluid chemistry can preserve evidence of biogenicity, erase such evidence completely or indeed can provide for post-depositional habitable conditions in the subsurface.  XRD / XRF data are critical to these determinations.

The CheMin instrument on the Mars Science Laboratory (MSL) Curiosity rover is the first XRD instrument flown in space. CheMin operates in transmission geometry with a Co X-ray source to minimize fluorescence from iron. Diffracted photons are collected with an energy-sensitive charge-coupled device (CCD). The position of the diffracted photons provides structural information for minerals, whereas the energy of sample-generated X-ray fluorescence photons provides elemental information, though these XRF data are qualitative. Mineralogical data from the CheMin XRD identified the three circumstances above: habitable depositional environments (e.g., Yellowknife Bay), habitable subsurface/diagenetic environments (e.g., throughout the Murray formation), and diagenetic conditions that may destroy evidence of habitability (e.g., oxidative and acidic environments at Vera Rubin ridge).

Technological advances in X-ray technology and lessons learned from the operation of CheMin on Mars have resulted in a next-generation XRD/XRF, called CheMinX. Replacement of CheMin’s CCD with an array of hybrid pixel detectors and improvements in focusing optics dramatically decrease analysis time (15 minutes vs. 22 hours for MSL-CheMin) and result in a better angular resolution (0.18 vs. 0.30 °2θ for MSL-CheMin). This increased resolution improves mineral detection, including discrimination between types of pyroxenes, which is not possible with MSL-CheMin data. The hybrid pixel detectors do not require cooling like the MSL-CheMin CCD, therefore reducing the power needed to operate CheMinX. CheMinX has a silicon-drift detector (SDD) to measure fluoresced photons, enabling the quantification of major, minor, and some trace elements via XRF. XRD/XRF data are collected simultaneously in CheMinX, obviating the need for multiple compositional instruments. The CheMinX design also improves upon MSL-CheMin’s sample handling. Instead of sample cells on wheel, which are often not reusable and add complexity in commanding the instrument, CheMinX has single-use cells in a cartridge/dispenser configuration. Because of these improvements, CheMinX is an ideal instrument for Discovery-class life-detection missions, including Mars Life Explorer that was recommended for development in the Planetary Science and Astrobiology Decadal Survey 2023-2032.

Session Description:

New In Situ Technologies for Searching for Signatures of Life on Mars

In situ analysis of biosignatures on Mars remains a challenge. Nevertheless, over the past two decades, substantial progress has been made on in situ technologies for life signature detection, spanning sample acquisition, processing, and analysis. Progress has resulted from field astrobiology studies of Mars analog sites on Earth as well as traditional flight opportunities (planetary, ISS, and cube-sats) and commercial activities, resulting in new miniaturized instrumentation with improved in situ science capabilities. We encourage contributions that cover a broad development spectrum, from early concepts to laboratory prototypes, modified flight-proven instruments and new mission architectures. We are particularly interested in contributions that address detection challenges using innovative approaches. With human exploration of Mars on the horizon, search for biosignatures and extant life on Mars has become a critical and timely endeavor.