## Slide 1
Medical Ceramic Oxygen Generator (M-COG)
Technology description presented to:
CSA Group Technical Subcommittee Z251.20 – Medical Gas Systems
May 6, 2024
John Graf
Technology Development Lead for Life Support
NASA Johnson Space Center, Houston TX USA
john.c.graf@nasa.gov
Trade names and trademarks are used in this report for identification only. Their usage does not constitute an official endorsement, either expressed or implied, by the National Aeronautics and Space Administration

## Slide 2
Applications for a Ceramic Oxygen Generator (COG)
Space Suit
O2 Tank Recharge
Ultra High Purity (UHP) O2
for analytical chemistry
Medical O2
In Challenging Environments
Key Requirements:
Purity: >99.99%
Pressure: >3000 psig
Safety: (no H2 hazards)
Key Requirements:
Purity: >99.999%
Price: < $500 / cylinder
Key Requirements:
Purity: >93% (prefer >99.0%)
Service Life: > 10 years
Operates in hot, humid environments
<1 hr Maintenance per year

## Slide 3: Attributes of Ideal Medical Oxygen Generator / M-COG Attributes
Reliability
Service Life
Training
Maintenance
O2 Purity
O2 Pressure
Temp. range
Humidity range
Elevation range
Power
Cost
MTBF >5yrs
>10 years
< 1 hr
<1 hr per year
>99.0%
>5 Bar
-10 to 50C
0 to 99.5%
0 to 4000 meters
<50 Watts/lpm
Less than SOA
one moving part, that part has 7 year MTBF)
>99.9%, > 5 bar after 10 years (35% reduction in flow)
Training needs: 1 hr instructional video
1 hr to replace blower & thermocouple every 2 years
>99.99% (under all operating conditions)
>13 Bar
-10 to 50C (limited by electronics)
0 to 99.5% (humidity no affect on purity, pressure, power)
0 to 4000 meters (demonstrated at 3300 meters)
Ideal System
M-COG
In Work <100 Watts/lpm now, <50 Watts/lpm with Gen2 wafer
In Work with mature manufacturing, equipment cost < $1000/lpm
Attribute

## Slide 5
The Two Main Components of an M-COG System: Oxygen Engine & Power Supply-Control
Power Supply & Control Assembly
(Ambient Temperature)
Oxygen Engine
(High Temperature)
Cell Stack Power Supply
Heater Power Supply
Blower Power Supply
Instrumentation
System
Controller
Oxygen Engine
(High Temperature)
Power Supply & Control Assy
(Ambient Temperature)

## Slide 6
Ion Transport Membranes

## Slide 7
Ion Transport Membranes

## Slide 8
M-COG Wafers

## Slide 9
M-COG Cell Stacks (components)

## Slide 10
Photograph of M-COG Cell Stack
Dimensions:
7.5 X 15 X 15 cm

## Slide 11
M-COG Boxed Cell Stack (exploded view)

## Slide 12
Photograph of M-COG Boxed Cell Stacks

## Slide 13
M-COG Process Schematic

## Slide 14: The COG power Dilemma (2012). Either too much cell stack power, or too much heater power.

## Slide 15
M-COG Process Schematic (4 cell stacks in series)

## Slide 16
M-COG System Prototype (referred to as Chassis 2)
Purity:
Production Rate:
Power Use:
Pressure:
99.995%
4.5 lpm for one stack (total O2 production 12.5 lpm)
400 Watts for 4.5 lpm (heaters were turned off, cell stack power only)
80 psig delivery pressure at WSTF (200 psig test pressure at AmOx)

## Slide 17
M-COG System Two-Stage Prototype (3000 psig delivery)
Simplified Schematic
1st Stage
Cell Stack
2nd Stage
Cell Stack

## Slide 18
A Comparison of M-COG Wafers to Bottle Filler Wafers

## Slide 19
Two-Stage Prototype Testing: October 2023
1st Stage
2nd Stage
2nd Stage
Cross section
2nd Stage
Interior, Top View
Key Measurement
Oxygen Pumping into pressure vessel with 3000 psig pressure

## Slide 20
M-COG System Prototype (referred to as Cross Flow Cylinder)

## Slide 21
Top view
section
Side view
section
1 cell stack
generator
(0-4 lpm)
Designed
Built
Tested
4 cell stack
generator (0-16 lpm)
Most Energy Efficient
Concept Only
2 cell stack
generator
(0-8 lpm)
Designed
Not-Yet-Built
1 prototype
Due June 2024
O2 Engine Configurations

## Slide 22
2024 Oxygen Engine (1 cell stack with 30 wafers)

## Slide 23
AmOx Key Measurements
Run time:
>2500 hours in 2023
5800 hours in 2024
Performance is nominal throughout the test.
System survived loss of facility power, unplanned system shut-down
Crossflow Cylinder Testing: Summer 2023 & Winter 2024

## Slide 24
Supply Aviator’s Breathing Oxygen (ABO) for US Military
In operation 24/365 for over 24-months
US Marine base in Afghanistan
System used by US Marines, operating a forward operating base in Afghanistan
Ceramic Oxygen Generator in Afghanistan: 2011 - 2013
Shipping crate after return from Afghanistan
Chemical Oxygen Generator operated nominally after return from Afghanistan
Housing for oxygen generator system
O2 Product purity was always >99.9%

## Slide 25
Aviator Breathing Oxygen (ABO)
Fill 10 full-size cylinders per week.
Fill 3 aircraft bottles per day
Deployed to Afghanistan to support US Troops
High Pressure Booster
(2 units)
SEOS Oxygen Engine
(2 of possible 3)
Oxygen
Receiver
High-pressure oxygen storage
Ceramic Oxygen Generator had a high pressure tank fill station
  Mechanical Compressors needed seal repair after 1 year
  Ceramic Oxygen Generator operated 24/7 for >2 years with perfect reliability

## Slide 26
Long Duration Wafer Testing
40 wafers tested >10,000 hrs (1.1 year)
7 wafers tested >40,000 hrs (4.5 yreas)
1 wafer tested >55,000 hrs (6.4 years)
Projected aging trends:
Slow, steady, gradual decrease in production rate (if system power is maintained). Decrease in production rate after 10 years continuous use: <30%

## Slide 27
M-COG Oxygen Contains Trace Amounts of Water Vapor
All available evidence suggests that water vapor is permeating through the seals of washers that separate wafers in the cell stack
  Delivery pressure trends
  Production rate trends
  Process air humidity trends
  Operating temperature trends
Water Vapor levels are affected by production rates
  At higher production rates, ppm H2O is lower
  At lower production rates, ppmv H2O is higher
The highest ppmv H2O measured at 3 lpm production rate: 58 ppmv
The highest ppmv H2O measured at 1.4 lpm production rate: 81 ppmv

## Slide 28
Regulator must remain
>11.9 C under fast flow
conditions (25 slpm)
To prevent icing
Max Pressure:
Max ppmv:
Dew point at
67 ppmv/3000 psig:
3000 psig
67 ppmv
11.9 C
O2
Medical O2
Humidifier
Trace Water Vapor in Compressed Gas: Medical O2

## Slide 29
Regulator must remain
> 5.7 C under fast flow
conditions (300 slpm)
to prevent icing. Two
stage regulators are
generally used
Max Pressure:
H2O max ppmv:
Dew point at
24 ppmv/5500 psig:
5500 psig
24 ppmv
5.7 C
SCBA Emergency Respirator
Air
Trace Water Vapor in Compressed Gas: SCBA

## Slide 30
Regulator must remain
> -18 C under fast flow
conditions (300 slpm)
to prevent icing. Two stage regulators generally used
Max Pressure:
H2O max ppmv:
Dew point at
6 ppmv/3000 psig:
3000 psig
6 ppmv
-18 C
Aviator’s Breathing Oxygen
O2
Trace Water Vapor in Compressed Gas: Aviator’s Breathing Oxygen

## Slide 31
Max Pressure:
H2O max ppmv:
Dew point at
250 ppmv/200 psig:
200 psig
250 ppmv
-6 C
Low Pressure M-COG (proposed)
Trace Water Vapor in Compressed Gas: M-COG (proposed)

## Slide 32
Max Pressure:
H2O max ppmv:
Dew point at
6 ppmv/3000 psig:
3000 psig
6 ppmv
-18 C
High Pressure M-COG (proposed)
Trace Water Vapor in Compressed Gas: High Pressure M-COG (proposed)

## Slide 33
Pressure
(psig)
Application
Max water
(ppmv)
Line Pressure
Dewpoint (C)
Gas
Notes / Comments
Med O2
Med O2
SCBA
SCBA
Aviator
M-COG (Low-P)
M-COG (Hi-P)
3000
2000
5500
3000
3000
200
3000
67
67
24
24
6
250
6
+ 11.9
- 5.9
+ 5.7
- 2.4
- 18.0
- 6.0
- 18.0
Ambient temps are common
2000 psi is a common use case
5500 psi for 45 min rating
3000 psi is a common use case
Cold temps likely in aviation
Receiver tanks rated to 200 psig
Actual water vapor <1 ppm
O2
O2
Air
Air
O2
O2
O2
Trace Water Vapor in Compressed Gas: Summary Table

## Slide 34
Oxygen Purity Can Be Verified by Inspection of Delivery Pressure
Ceramic Ion Transport membranes are fundamentally selective to oxygen. Only oxygen ions can pass through the electrolyte layer
  Electrolyte layer is gas impermeable. At 200 psig, the leak rate for each wafer is <0.013 sccm
M-COG process air is moved using a blower. The maximum pressure the blower can achieve is <1 psig.
If the gaseous product delivered by M-COG has a pressure >1psig, the pressurized gas was produced by the ion transport membrane.
Municipal Drinking Water Supply uses pressurized “outleakage” methods to verify purity using delivery pressure.
Delivery pressure of >1 psi guarantees the oxygen purity is >99.5%
  >1 psig pressure verifies the source of the oxygen is M-COG ion transport membrane
  M-COG produced oxygen purity is always >99.5%
  Purity has been verified by NASA test
    M-COG system had a fracture at the ceramic/metal joint
    M-COG delivery pressure <0.5 psi
    M-COG purity 99.97% (nitrogen 240 ppmv)

## Slide 35
M-COG System Prototype Testing: August 2021
Purity was 99.97% under worst case in-leakage conditions: The M-COG gas delivery system was fractured and had no pressure integrity. Maximum delivery pressure was < 0.5 psi.

## Slide 36
Medical Oxygen Specifications
WHO International Pharmacopoeia Definition of Medical Oxygen
Oxygen:
Carbon Dioxide:
Carbon Monoxide:
Water Vapor:
Not less than 99.5% v/v
Not more than 300 ppm v/v
Not more than 5 ppm v/v
Not more than 67 ppm v/v
Recommended Specifications for M-COG Produced Medical Oxygen
Oxygen:
Carbon Dioxide:
Carbon Monoxide:
Water Vapor:
Not less than 99.5% v/v
Not more than 300 ppm v/v
Not more than 5 ppm v/v
Not more than 250 ppm for gases stored at pressures less than 15 Bar
Not more than 6 ppm for gases stored at pressures between 15 and 200 Bar
USP 29 Standards for Medical Oxygen
Oxygen:
Carbon Dioxide:
Carbon Monoxide:
Water Vapor:
Not less than 99.0% v/v
Not more than 300 ppm v/v
Not more than 10 ppm v/v
n/a (no specification for water vapor)
Note:
Oxygen that is produced by the air liquefaction process is exempt from the requirements of the tests for carbon dioxide and carbon monoxide

## Slide 37
Demonstration System Concept:
  15 m3/hr O2 (250 lpm)
  200 kg/day steam
  2500 l/day hot water
  25kW continuous power
Electronics
Steam &
Hot water
O2
Engines
Crate Packaging Plan
Heat Exchanger
O2
Storage

## Slide 38
Summary
M-COG performance is insensitive to temperature.
M-COG performance is insensitive to elevation/atmospheric pressure.
M-COG performance is not compromised by humidity.
M-COG can use any air (provided it is safe to breathe) without filtration or preconditioning.
M-COG can operate at small scale
M-COG is easy to ship
M-COG does not need a mechanical compressor.
M-COG is energy efficient.
Compared to PSA, M-COG requires less maintenance.
M-COG oxygen is 99.9% pure
A two-stage version of M-COG can deliver high-pressure oxygen (capable of cylinder recharge)
In some operating environments, M-COG can be the most reliable and method of oxygen supply.