## Slide 1: Specific Heat Capacity Measurement of Bi-Carbide Fuels for Nuclear Thermal Propulsion
Jason Reynolds1, Jhonathan Rosales1, Martin Volz1, Brian Taylor1, Jamelle Williams1, Arne Cröll2, Travis Knight3, Ted Besmann3, Juliano Schorne-Pinto3, Ronnie Booth3, Matt Phipps3, Robert Demuth3
1NASA George C. Marshall Space Flight Center
2University of Alabama in Huntsville
3University of South Carolina

## Slide 2
Instrumentation
Heat capacity measurements: DSC 404 F3 Pegasus® - Netzsch
Reducing O2 contamination: SETNAG (purifier) + sacrificial Zr ring (inside DSC)
Methodology
ASTM - 2018 - E1269-11 Standard Test Method for Determining Specific Heat Capacity by DSC
“Classical” three-step method: 1) empty crucible, 2) sapphire standard, and 3) sample. Conditions:
Temperature range: 100-1200/1360 °C
Heating rate: 10 K/min
Atmosphere: He (Oxygen levels down to ~10-23 atm)
Crucibles: Pt-Rh
Verification of measurement precision using sapphire:
Type B error: ~1-2%
Maximum inaccuracy: 0.7%

## Slide 3
Estimation of heat capacity for the mixtures
Neumann Kopp (NK) is an additive semiempirical approach to obtain the Cp(T), also known as linear rule of mixtures (ROMs)
Example: 32UC-68ZrC = 0.32Cp(T)UC + 0.68Cp(T)ZrC
Chart: Uranium monocarbide (UC)
Chart: Zirconium monocarbide (ZrC)
Chart: Estimated: 32UC-68ZrC

## Slide 4
Uncorrected values + NK estimations
Corrected sample#2 + NK
Outcomes:
Results from sample#1 were discarded because of the oxidation issues due to the long DSC runs.
It seems that (32UC-68ZrC)solid-solution converges to NK/ROM values at high-temperature.
Samples after DSC:
If the sample was fully oxidized:
(32UC-68ZrC)solid-solution = ZrO2 + UO2
The mass would change +13.3%
Sample#1 32UC-68ZrC:
Initial mass = 83.95mg (N=5)
Final mass = 84.63mg (N=5)
Variation = 0.68mg (+0.8%), equivalent to oxidation of ~ 6.1% of the monocarbides
Sample#2 32UC-68ZrC:
Initial mass = 85.40mg (N=5)
Final mass = 85.53mg (N=5)
Variation = 0.13mg (+0.15%), equivalent to oxidation of ~ 1.1% of the monocarbides
Preliminary results (32UC-68ZrC)
#1
#2

## Slide 5
Outcomes:
Unexpectedly high Cp* in comparison to NK/ROM
New tests will be required to statistically validate this preliminary result (two runs, single sample), varying sample size and shape**. Powder samples would be preferable, but there is a high risk of oxidation during the DSC run.
Other DSCs might be used (STA at USC or Setsys at NASA)
Sample after DSC:
Preliminary results (10UC-90ZrC)
5
Sample 10UC-90ZrC:
Square shape with a reduced contact area with the detector compared to the preceding sample (32UC-68ZrC)
Initial mass = 63.94mg (N=5)
Final mass = 63.93mg (N=5)
Variation = -0.01mg (-0.02%), no oxidization could be quantified
1 Synthesis, thermal conductivity, and hydrogen compatibility of a high melt point solid solution uranium carbide, (U0.2Zr0.8)C, Nucl. Mater. Energy, v 33, 101290
2 Thermal and electrical properties of a high entropy carbide (Ta, Hf, Nb, Zr) at elevated temperatures, J. Am. Ceram. Soc., v 105(6), p. 4426-4434
*Carbide mixtures, according to Kardoulaki et al. (2022)1 and Schwind et al. (2022)2, diverge from NK/ROM at low temperature, but eventually converge at high temperatures to the ideal mixture behavior.
**Heat flux DSC results can be very dependent of the sample format and thermal contact with the 2D detector

## Slide 6
UC-ZrC solid solution
1) Comparing with UC and ZrC:
2) Experimental vs Neumann Kopp:

## Slide 7: Acknowledgements
The current work was completed at NASA - Marshall Space Flight Center and supported by NASA’s Space Technology Mission Directorate (STMD) through the Space Nuclear Propulsion (SNP) project
This work was achievable with the successful powder synthesis by Eri Kardoulaki and Ken McClellan at Los Alamos National Laboratory and expertise from Travis Knight and Ted Besmann’s team at the University of South Carolina.