## Slide 1
CLDP Draft Utilization Requirements

## Slide 2: DRAFT Utilization Requirements
Conditioned Stowage
Gloveboxes
Microscopy
Plant Research
Pressurized Payload Resources
Unpressurized Payload Resources
Data Management and Comm
Lab Equipment and Supplies
  Biological and Physical Centrifuge
  Ultrasound
  Laboratory Optics Bench
  Sample Collection Equipment
  Basic Utilization Consumables
Environmental Monitoring
Other Applicable Requirements
All requirements are considered draft and subject to change prior to final RFP release.
All requirements are considered draft and subject to change prior to final RFP release.
Requirements from CLDP-REQ-1130 are shown for feedback
Notes: This package contains the draft utilization requirements
Text in orange block will replace the CUI banner at the bottom of each chart in the package.

## Slide 3: DRAFT Conditioned Stowage Requirements
All requirements are considered draft and subject to change prior to final RFP release.

## Slide 4: DRAFT Cold Stowage Capability
All requirements are considered draft and subject to change prior to final RFP release.
 | Min. Temperature Range | Min. Volume - Destination | Min. Volume - Transportation
Refrigerated | +3C to +22C | 100L | 50L
Frozen | -100C to -3C | 350L | 50L
Cryogenic | -160C to -100C | 5L | Optional 5L**
Conditioned stowage facilities provide the capability to preserve collected science in appropriate temperature environments and transfer safely back to Principal Investigators.
**Currently not an ISS capability
Notes: ISS Capabilities
Glacier (8.38L) = -95C to +4C (transportation) and -160C to +4C (ISS)
Iceberg (18.9L) = -95C to +4C
Polar (10.99L) = -95C to +10C
MERLIN (12.45L) = +4C to +40C (transportation) and -20C to +49C (ISS)
MELFI (approx. 300L per unit) = +2C, -35C, -95C
Rapid Freeze (CryoChiller) = 25ml from +22C to -40C in 5 minutes and to -120C in 10 minutes
We have not included Passive CS Hardware in the requirements
Why +/-3C?
To allow for temperature stability range around the phase change region of 0 deg C
Volumes (data from last 15 flights in backup)
Cryogenic = 5L accommodates anticipated BPS needs and is less than current Glacier capability
Frozen = Largest use on ISS with average of 625L in the past 15 flights
Refrigerated = Second largest use on ISS with average of 250L in the past 15 flights
Transportation allows for 2-3 Polar equivalents and 2-3 Merlin equivalents [currently average 7ish units in travel]
Discuss cryogenic as an option or not available on ISS but would open additional opportunity for plasma science

## Slide 5: DRAFT Incubation Capability
All requirements are considered draft and subject to change prior to final RFP release.
 | Min. Temperature Range | Min. Volume - Destination | Min. Volume - Transportation
Incubation | +6C to +49C | 20L | 20L
Other Environmental Control and Monitoring
  Oxygen, humidity, and light
Surface Disinfection
  Disinfect all surfaces within the incubator from released molecular, biological, and chemical contaminates to levels safe for crew exposure
Notes: ISS Capabilities
MERLIN (12.45L) = +4C to +40C (transportation) and -20C to +49C (ISS)
SABL (23L) = -5C to +43C

## Slide 6: DRAFT Conditioned Stowage Form Factors
All requirements are considered draft and subject to change prior to final RFP release.
Representative Form Factors | Cryogenic | Frozen | Refrigeration | Incubation
Cryovials 49mm height and 13mm diameter | x | x |  | 
Cell Culture Chambers 15cm x 15cm x 15cm | x | x | x | x
Small Re-Sealable Storage Bags 12cm x 15cm x 3cm | x | x | x | x
Large Re-Sealable Storage Bags 16cm x 16cm x 3cm | x | x | x | x
Large Experiment Modules 35cm x 30cm x 25cm |  | x | x | 
Notes: Data provided by JSC Cold Stowage as the most flown form factors for the current units

## Slide 7: Other DRAFT Conditioned Stowage Capabilities
All requirements are considered draft and subject to change prior to final RFP release.
Rapid Freeze
  Rapidly freeze samples of at least 25ml from +22°C to -40°C in 5 minutes and to -120°C in 10 minutes while within one level of containment
Failure Tolerance
  Single failure tolerant to preserve science
Configurable Setpoints
  at least two reconfigurable setpoints within the given ranges
Sample Temperature Stability
  within +/- 3°C or better when sample is at equilibrium
Notes: Rapid Freeze = current ISS capability
Failure Tolerance = could be designed into systems or keep available spares
Configurable Setpoints = allows use of multiple temperatures simultaneously within the given ranges
Temperature Stability = JSC engineering states +/- 3C as the current fluctuation we see in sample temp stability

## Slide 8: DRAFT Glovebox Requirements
All requirements are considered draft and subject to change prior to final RFP release.

## Slide 9: General – DRAFT Glovebox Requirements
All requirements are considered draft and subject to change prior to final RFP release.
Gloveboxes provide the capability to accomplish a variety of NASA investigations including biotechnology, combustion science, fluid physics, fundamental physics, and material science
Glovebox
Remote-tended Experiments
Crew-tended Experiments
Health and Status Monitoring
Real-Time Imagery
450L Volume
Notes: NASA anticipates pursing enough science throughput to utilize potentially two gloveboxes
LSG = 450L = 16ft^3
MSG = 255L = 9ft^3

## Slide 10: DRAFT Glovebox Containment
All requirements are considered draft and subject to change prior to final RFP release.
Two Levels of Containment
Keeps the crew safe from hazardous materials
Historically, has been accomplished by (1) a fully sealed physical barrier and (2) differential pressure, air circulation, and air filtration
Cleanliness
Maintain glovebox work volume to at least a Class 100,000 cleanliness level per Federal Standard Airborne Particulate Cleanliness Classes in Cleanrooms and Clean Zones (FED-STD-209E)
Air Filtration
Filter biological and chemical particulates of at least 0.3 microns
Notes: Current ISS glovebox capabilities:
Cleanliness
  Class 100K
Containment
  Two levels (1 is fully seal physical barrier and 2 is the list below)
Differential Pressure
  at least 1.3 mB to 7mB relative to the ambient destination pressure
Air Circulation
  Max airflow rate of 1200 l/min and max velocity of 0.44 m/s at the center line of the glovebox
Air Filtration
  Filter biological and chemical particulates of at least 0.3 microns – HEPA filters

## Slide 11: DRAFT Glovebox Requirements
All requirements are considered draft and subject to change prior to final RFP release.
Min. Glovebox Payload Accommodations
Glovebox Payloads | 
Mechanical Attachment | At least 2 reconfigurable attachment points
Power Connections | 120VAC, 28VDC, 12VDC, and 5VDC
Power to Payloads | 1kW
Heat Rejection | 1kW
Vacuum Exhaust | 1x10^-3 torr
Vacuum Resource | 1x10^-3 torr
Additional Glovebox Payload Accommodations
  Cabling feed-through with 5cm cross section
  Payload circuit protection to prevent excess and unsafe loads of power
  Capability to access water for use within the glovebox
  Capability to access up to 5.43kg/hr of nitrogen gas for use within the glovebox
  Data connection within the glovebox work volume
Notes: MSG
120VDC, 28VDC, +/-12VDC, and 5VDC
800W coldplate + 200W air
Nitrogen, VES, and VRS
LSG
120VAC, 28VDC
350W

## Slide 12: DRAFT Glovebox Requirements
All requirements are considered draft and subject to change prior to final RFP release.
Glovebox Crew Accommodations | 
Glove Sizing | Variety of sizes in accordance with NASA-STD-3001 V2
Decontamination | Protects crew from exposure to released molecular, biological, and chemical contaminates
Illumination | At least 400-1000 lumens at 200mm above surface of work bench
Accessibility | Multiple entry points into glovebox from a variety of angles and directions
Multi-Crew Access | At least 2 crew working in 1 glovebox simultaneously with both sets of hands
Reachability | Ability to reach every area of the internal work volume during operations
Workbench Access | Access to the largest surface of the work volume for payload setup and teardown
Notes: Glove Sizing – To accommodate crew members of various sizes
Decontamination – To protect the crew (could be accomplished by UV lights, cleaning procedures, etc.)
Illumination – Current ISS capability, meet or exceed
Accessibility – Capability for an entry point from more than 1 angles (above, below, in front, on side, etc.)
Multi-Crew – At least 2 crew should be able to work in glovebox together
Reachability – ISS lessons learned where crew member with any size gloves should be able to reach every part of the internal work volume
Workbench Access – For setup and breakdown of payloads (i.e. detachable front panel)

## Slide 13: DRAFT Microscope Requirements
All requirements are considered draft and subject to change prior to final RFP release.
Notes: Need to update to reflect biological science microscope vs physical science microscope

## Slide 14: DRAFT Microscope Requirements
All requirements are considered draft and subject to change prior to final RFP release.
Microscopes provide the capability to biological science and physical science microscopy investigations for imaging samples difficult to see with the human eye
Microscope
Crew-tended Operations
Remote-Tended Operations
Data Management – store, process, and downlink raw data and imagery
Temperature Control – at least 0C and 85C with a +/- 0.1C stability
Notes: Containment of samples will be work into the sample modules

## Slide 15: DRAFT Microscope Sample Accommodations
All requirements are considered draft and subject to change prior to final RFP release.
Microscope Sample Accommodations | 
Sample Modules | Between 60x200x4 mm and 60x200x250 mm
Microscope Slides and Chamber Slides | Up to 76x26 mm
Petri Dishes | Between 35 mm and 90 mm
Cell Culture Plates | 96-well and 384-well
Microtiter Plates | Up to 85x128 mm
Sample Module Adjustability
Adjustable in x and y directions
10 µm increments
Between 10 µm/sec and 1mm/s

## Slide 16: DRAFT Microscope Lenses Suite
All requirements are considered draft and subject to change prior to final RFP release.
Various Magnifications and Objective Lenses
Dry lenses
Long working distance objectives
Water immersion lenses with a temperature control between 0C and 40C
Oil immersion lenses with a temperature control between 0 and 40C
Bertrand lenses
Notes: Berek Compensator – too much detail

## Slide 17: DRAFT Microscope Imaging Requirements
All requirements are considered draft and subject to change prior to final RFP release.
General Microscope Imaging Requirements | Parameters
Imagery | Real-time and recorded imagery between 1 image/hour and 50 images/second
Resolution | Centroid resolution of at least 20nm x 20 nm +/-2 nm x 100 nm
Confocal Microscopy | Imaging static and dynamic fluorescent particles ranging from at least 0.01 to 5.0 microns in size and with fluorescent signals in the 450 to 900 nm range
Biological Science Imaging Requirements | Parameters
Biological Science Imaging | Real-time and recorded imagery for the following operational use-cases a. automated high throughput imaging of 96-well plates b. live and fixed cell imaging c. imaging small animal tissue samples d. automated imaging and stitching of a large area across multiple fields of view

## Slide 18: DRAFT Microscope Illumination Requirements
All requirements are considered draft and subject to change prior to final RFP release.
Illumination Requirements | Parameters
Illumination Intensity | a. Trans-illumination with maximum intensity of 3W/3mm and 10% uniformity b. Epi-illumination with maximum intensity of 3W/3mm and 10% uniformity c. Kohler illumination
Illumination Spectrum | a. White light source with spectral uniformity of +/- 10% over the wavelength range of 350 to 780 nm b. Red light source with a spectral uniformity of +/- 10% over the wavelength range of 532 to 650 nm c. Optically coupled laser light source with output adjustable between 5 and 10%, accuracy and precision of </= 1%, stability of 0.1%, linear polarization of >/= 500:1, and collimated at the sample d. Bright field imaging e. Dark field imaging
Polarizer Analyzer | Rotatable 360° in 1° increments with minimum extinction ratio of 5000:1
DIC Prisms | Differential Interference Contrast (DIC) imagery
Fluorescence Light Source | a. Confocal fluorescence light sources (at least 445 nm, 488nm, and 561nm) b. Wide-field fluorescence light source (at least 350-700nm)

## Slide 19: DRAFT Microscope Requirements for Physical Science Investigations
All requirements are considered draft and subject to change prior to final RFP release.
Physical Science Microscopy Requirements | Parameters
Microscope Power Availability | +/- 12 VDC and +/-5 VDC
Temperature Gradient | Sample temperature gradient up to 10C along the test cell length
Acoustic Force Control | Acoustic force between at least 0 and 50 dB with a variable frequency of 0 to 10 KHz
Centroid Resolution | At least 20nm x 20 nm +/-2 nm x 100 nm
Homogenization | Mix test samples and monitor progress using a magnetic stir bar
Optional Optical Tweezer Requirements* | Parameters
Microscope Optical Tweezers | Static and dynamic trapping for single and multiple particle manipulation
Microscope Particle Displacement Velocity | Minimum particle displacement velocity of 0.1 µm over 20 seconds Maximum particle displacement velocity of 100 µm over 0.1 seconds
*Currently not an ISS capability
Notes: Optical Tweezers
Utilized for much of the current soft matter research
Would provide the capability to enhance crystal growth experiments in LEO

## Slide 20: DRAFT Plant Research Requirements
All requirements are considered draft and subject to change prior to final RFP release.

## Slide 21: DRAFT Environmentally Controlled Habitat
All requirements are considered draft and subject to change prior to final RFP release.
Min. Environmentally Controlled Plant Habitat Accommodations | 
Volume | Growth Area = 1708 cm2 Growth Height = 43 cm Root Area = 1853 cm2 Root Height = 5 cm
Illumination | Controllable photon flux of 0-1000 µmol/m2s (over the PAR 400-750 nm wave band) with the ability to add or subtract the following: a. red (0-600 µmol/m2s at 630-660 nm ±10 nm) b. blue (0-400 µmol/m2s at 400-500 nm ±10 nm) c. green (0-100 µmol/m2s at 525 nm ±10 nm) d. broad spectrum white (0-600 µmol/m2s at 400-700 nm) e. far-red (0-50 µmol/m2s at 730-750 nm ±10 nm)
Nutrient Delivery | 2L/m2/day to 5L/m2/day
Data Management | Record and transfer payload data, science data, and imagery data
Environmental Measurements | Measure, record, and provide light and illumination levels and surface temperature
Environmental Controls | a. Temperature between at least 16-30°C b. Humidity between at least 40-90% c. Carbon Dioxide between at least 400 to 5000 ppm +/- 50ppm
An environmentally controlled and enclosed plant habitat allows plant bioscience investigations in a microgravity environment but with enhanced controls over other environmental factors.
Notes: Photon flux = measure of light intensity
PAR = Photosynthetic Active Radiation

## Slide 22: DRAFT Open-Air Plant Habitat
All requirements are considered draft and subject to change prior to final RFP release.
Min. Open-Air Plant Habitat Accommodations | 
Volume | Growth Area = 24 cm2 Growth Height = 45 cm Root Area = 16 cm2 Root Height = 5 cm
Illumination | LED illumination with adjustable red, blue, and green wavelengths
Nutrient Delivery | 250 mL/day to 2 L/day
Data Management | Record and transfer payload data, science data, and imagery data
Environmental Measurements | Measure, record, and provide temperature, humidity, and CO2 levels
An open-air plant habitat allows the pursuit of plant bioscience and observe how plants use nominal cabin environmental factors to grow and produce in microgravity.

## Slide 23: DRAFT Pressurized Payload Requirements
All requirements are considered draft and subject to change prior to final RFP release.

## Slide 24: DRAFT Internal Payload Accommodations
All requirements are considered draft and subject to change prior to final RFP release.
Min. Per Payload Accommodations | 
Mechanical Attachment | Yes
Power | 500W
Voltage | 28VDC
Heat Rejection | 500W
Min. Shared Resource Accommodations | 
Vacuum Exhaust | 1x10-3 torr
Vacuum Resource | 1x10-3 torr
Gasses | N2 and CO2
Other | Power Inverter and Converter
Pressurized payload accommodations provide the capability to perform science investigations in microgravity with proposed resources in a reconfigurable multi-host facility
Notes: Based on current EXPRESS accommodation capabilities

## Slide 25: DRAFT Unpressurized Payload Requirements
All requirements are considered draft and subject to change prior to final RFP release.

## Slide 26: DRAFT External Payload Accommodations per Payload Site
All requirements are considered draft and subject to change prior to final RFP release.
Min. Unpressurized Payload Accommodations | 
Mechanical Attachment | Yes
Mass | 500kg
Volume | 1.5m x 1.5m x 3m
Power on Destination | 1kW operational 300W survival
Power during Translation | 300W for 6 hours survival during translation
Voltage | 120 VDC
Heat Rejection | 1kW
Installation and Removal | Yes
Other | Dual-Power Channel
***Field of View will be addressed in Concept of Operations
Unpressurized payload accommodations provide the capability to perform Earth Science, Heliophysics, Planetary Science, and Astrophysics investigations by utilizing viewpoints around the destination.
Notes: Based on current standard JEM-EF accommodations

## Slide 27: DRAFT Data Management and Comm Requirements
All requirements are considered draft and subject to change prior to final RFP release.

## Slide 28: DRAFT Data Management
All requirements are considered draft and subject to change prior to final RFP release.
Min. Data Rates and Storage | 
Downlink Rate | 1 Gbps
Uplink Rate | 100 Mbps
Transfer Rate | 1 Gbps
Data Storage | 100 TB
Data Latency | <1 second
Rates and Storage
Additional Data Requirements
Process commands, handle generated data, and execute autonomous functions
Monitor and record time-synced real-time payload telemetry, health, and status
Record data during loss of signal
Host payload specific software applications
Secure, private, and encrypted storage, transmission, and distribution of payload data between (1) the Destination and Ground Control centers and (2) Destination and payload data owners

## Slide 29: DRAFT Lab Equipment and Supplies Requirements
All requirements are considered draft and subject to change prior to final RFP release.

## Slide 30: DRAFT - Biological Centrifuge
All requirements are considered draft and subject to change prior to final RFP release.
Performance Parameters | 
Timing | 1 minute to 99 minutes in 1-minute increments
Indefinite Runs | Allow for indefinite centrifugation run times
Protocols | Uplink programmable centrifugation protocols
Multi-Step Protocols | Uplink programmable multi-step protocols including holding/dwelling periods of varying RPM including zero
Force | Up to 3280 Relative Centrifugal force (RCF)
Speed | Configurable between 500 to 4000 RPM selectable in 100 RPM increments
Sample Sizes | 0.5, 1.5, 2.2, 3, 5, 7, 12, 15, and 50 ml.
Volume Size | Six 50ml vials
Safety Parameters | 
Sample Installation Alerts | Positive indication that samples are installed correctly
Protocol Completion Alerts | Positive indication that protocol has ended
Emergency Stop | Remotely and manually stop capability
Emergency Lid Release | Capability to retrieve the sample in case of power failure
Imbalanced Condition | Detection and forced automatic shut down in case of unbalanced condition
Controlled Acceleration and Deceleration | Capability to manually control the rotor angular acceleration and deceleration
Contamination Prevention | Prevent the cross-contamination of human and non-human samples
Provides the capability to separate substances of varying sample densities such as red blood cells, platelets, and plasma via centrifugal force applications

## Slide 31: DRAFT Physical Centrifuge
All requirements are considered draft and subject to change prior to final RFP release.
Performance Parameters | 
Simulated Gravity Control | At least 0.1g to 2g
Sample Accommodation | Average sample weight of 0.3kg
Volume | At least 390mm total volume Accommodate at least twelve modules of at least 800cc each
Environmental Controls | a. Temperature between 14C and 40C b. Humidity between 50% and 80% c. Carbon dioxide between 4.5% and 5.5%
Lighting | Visible lighting from at least 400-700nm and IR lighting
Imagery | Real-time imagery and video downlink
Provides the capability to utilize an environmentally controllable centrifuge to simulate up to 2g artificial gravity

## Slide 32: DRAFT Ultrasound (1/2)
All requirements are considered draft and subject to change prior to final RFP release.
Heart Scanning
Mass of the left and right ventricular Volume of each heart chamber Heart functions including but not limited to: a. Systolic and diastolic cardiac function b. Heart valve function c. Electrocardiogram data d. Respiratory trace data
Provides the capability to image, measure, and collect data on internal body structures of humans and animals for diagnostics and utilization
Bladder Scanning
Bladder Volume and Wall Thickness Bladder Functions Ureteral jets Dynamic motions
Vascular Scanning
Superficial and deep blood vessels Blood vessel structure and layers Vascular Blood Flow Electrocardiogram data Respiratory data Vascular Pathologies
Eye Scanning
Eye structure Ocular blood flow
Cerebral Scanning
Cerebral region scanning Cerebral blood flow
Bone Scanning
Cortical bone structures Detecting fractures
Abdominal Scanning
Soft tissue Small parts Compartments of the abdominal region Pathologies and fluid dynamics
Muscle Scanning
Muscle Pathologies Muscle Tissue Volume

## Slide 33: DRAFT Ultrasound (2/2)
All requirements are considered draft and subject to change prior to final RFP release.
Performance Parameters | 
Scanning Modalities | a. Two-Dimensional Scanning (READ Zoom, WRITE Zoom, and Harmonic Imaging) b. Doppler Scanning (Pulsed and Continuous Wave Spectral Doppler, Color Flow Doppler, Amplitude Doppler, Spectral Tissue Doppler, and Color 2-D Tissue Doppler) c. M-mode and Color Doppler M-mode Scanning d. Three-Dimensional Scanning (Panoramic and Real-time Volume) e. Tri-plexing between 2-D and Doppler Modalities f. Dual-image Display for Comparison Scanning
Imagery | Privatized real-time imagery and video downlink
Data Collection | Capture, store, and downlink still image data, dynamic image loops, device native data, diagnostic data, and user command data in a nonproprietary format.
Display Characteristics | a. Repositionable by operator b. Viewing angle of at least 135 degrees c. Minimum screen size of at least 39.624 cm d. Minimum resolution of at least 1920x1080 pixels e. Color depth of at least 8 bits f. Display of timestamp in hours, minutes, and seconds g. Display of instrument parameter information (i.e. Mechanical Index and Thermal Index)
Control Panel | a. Repositionable by operator b. Fine pointer control for user input c. Alphanumeric data entry for user input

## Slide 34: DRAFT Laboratory Optics Bench
All requirements are considered draft and subject to change prior to final RFP release.
Min. Lab Optics Bench Accommodations | 
Mechanical Attachment | Configurable, movable, and adjustable
Voltage | 120VAC, 28VDC, 12VDC, and 5VDC
Data Acquisition | Video and digital data acquisition
Optical Precision Alignment Surface | Precision Alignment
Stable Thermal Environment within +/- 5C | 
Provides the capability perform configurable payload utilization experiments on a laboratory optics bench
Notes: Data Acquisition
Video and digital data acquisition

## Slide 35: DRAFT Sample Collection Equipment
All requirements are considered draft and subject to change prior to final RFP release.
Human Sample Collections | 
Blood Collection Consumables | Gloves, butterfly needles, gauze, band aids, biocide wipes, dry wipes, tourniquets, sharps containers and biohazardous liner bags
Blood Collection Tubes | Variety of tubes including EDTA (4 ml, 6ml, 7ml), Heparin (6ml, 7ml, 10ml, Sodium), ACD (6ml, 8.5ml), CPT (4ml), Serum (7ml), and RNA (2.5 ml).
Urine Collection | Urine collection devices, urine sample bags for determining void volume, and capability for both male and female crewmembers
DNA Collection | Swabs, tubes for collecting and returning samples from crewmembers depending on experiment requirements.
Saliva Collection | Salivates and sample storage containers for dried, rolled, sublingual, and timed collections
Fecal Matter Collection | Sample containers and aliquot kits for collecting, processing, and returning
Human Sample Collection Labeling | Uniquely identifying and distinguishing labels to separate each type of samples and from other crewmembers

## Slide 36: DRAFT Laboratory Supplies and Equipment
All requirements are considered draft and subject to change prior to final RFP release.
Generic Laboratory Supplies and Equipment | 
Personal Protective Equipment | Nylon gloves, safety glasses/goggles, ear plugs, and face shields
General Consumables | Plastic storage bags (Ziplock), Kapton tape, Velcro, and water bags
Sample Homogenization Equipment | Agitators, mixers, and vortexes
Sample Handling Equipment | Test tubes, well plates, flasks, syringes, pipettes, and sieves/filters suitable
Mass Measurement | Small mass samples with high precision measuring Human sized samples
Laboratory Analysis Equipment |  |  | 
Fixation | Fluorescence Activated Cell Sorting | Gas Chromatography | Evaporator
RNA/DNA Sequencing | Cell Counting | Static and Dynamic Light Scattering | Distillation
Polymerase Chain Reaction | Tensile and Compression testing | Nuclear Magnetic Resonance | Laminar Flow Hood/Bench Equivalent
Spectrophotometry | Protein Analysis | pH Conductivity | Glass Stirred Reaction
High Performance Liquid Chromatography | High Performance Liquid Chromatography | Spectroscopy (UV, Visible, Fourier Transform Infrared, and Raman) | 

## Slide 37: DRAFT Environmental Monitoring Requirements
All requirements are considered draft and subject to change prior to final RFP release.

## Slide 38: DRAFT Environmental Monitoring
All requirements are considered draft and subject to change prior to final RFP release.
NASA utilization requires the capability to establish an environmental baseline and monitor disturbances to that baseline
Environmental Monitoring
Radiation
Atmosphere
Spacecraft Pointing and Attitude
Disturbance
Contamination
Electromagnetic Interference
Global Positioning
Microgravity

## Slide 39: DRAFT Microgravity Environment
The Destination shall provide the following microgravity acceleration performance for at least 50% of the internal payload locations for 50% of the increment in continuous time intervals of at least 30 days:
At the centers of the internal payload locations, a quasi-steady (<0.01 Hz) acceleration.
  Magnitude less than or equal to 1 micro-g
  Component perpendicular to the orbital average acceleration vector less than or equal to 0.2 micro-g
At the structural mounting interfaces to the internal payload locations
  A vibratory acceleration limit as defined in the figure below.
  A transient acceleration limit for individual transient disturbance sources less than or equal to 1000 micro-g per axis.
  An integrated transient acceleration limit for individual transient disturbance sources less than or equal to 10 micro-g seconds per axis over any 10 second interval.
All requirements are considered draft and subject to change prior to final RFP release.
Adopted from ISS microgravity levels in SSP 41000
Acceleration performance allows for support of a wide-range of microgravity science
Notes: Out of a joint board between Utilization, SE&I, and CCP, it was noted that a microgravity requirement to accomplish microgravity science was missing.
--- It is the recommendation of the UZ office to adopt the microgravity requirement from SSP 41000. This is the requirement for ISS microgravity.
--- This allows maximum support for a wide-range of microgravity science
--- We wanted to bring this requirement to this group to see if anyone has any issues, concerns, or comments to adding this requirement as stated
RMS = root mean square
Octave = unit of frequency level when the base of the logarithm is 2
--- Octave band is a frequency band where the highest frequency is twice the lowest frequency
--- 1/3 octave band is a frequency band whose upper band-edge is the lower band time the cube root of two
R.CLDS.534 is the Microgravity Monitoring requirement

## Slide 40: Other DRAFT Applicable Requirements
All requirements are considered draft and subject to change prior to final RFP release.

## Slide 41: DRAFT Time Critical Cargo
All requirements are considered draft and subject to change prior to final RFP release.
On-Orbit Cargo Removal
  Remove late-load cargo within 96 hours from ground handover
On-Orbit Cargo Loading
  Load cargo into vehicle less than 24 hours prior to vehicle departure from Destination
Pre-Launch Cargo Handling
  Install and maintain cargo within 24 hours of a scheduled launch
Ground Cargo Removal
  Transfer to USG representatives within 1 hour of landing
Hard Mounted Cargo (Best Practice)
  Accommodate powered and unpowered time critical cargo from pre-launch through transfer to Destination
Notes: Rapid Freeze = current ISS capability
Failure Tolerance = could be designed into systems or keep available spares
Configurable Setpoints = allows use of multiple temperatures simultaneously within the given ranges
Temperature Stability = JSC engineering states +/- 3C as the current fluctuation we see in sample temp stability

## Slide 42: DRAFT Privatized Data and Comm
All requirements are considered draft and subject to change prior to final RFP release.
Private Audio and Video
  Secure private audio and video communications between USG crew and USG ground control centers
Advanced Encryption Methods
  Secure hardware and software implementation for (1) decryption of commands of external origin and (2) encryption for commands with external destinations
    Excludes FTS commands for FIPS Publication 197, AES
Notes: Rapid Freeze = current ISS capability
Failure Tolerance = could be designed into systems or keep available spares
Configurable Setpoints = allows use of multiple temperatures simultaneously within the given ranges
Temperature Stability = JSC engineering states +/- 3C as the current fluctuation we see in sample temp stability

## Slide 43: DRAFT Human Factors Requirements (1/2)
All requirements are considered draft and subject to change prior to final RFP release.
Controls for Human Error
  Implement controls to prevent human error in maintenance, operation, and control; reduce likelihood; design recovery; limit negative effects of human errors
Use of Color
  Red = Emergency, Warning, Failed
  Yellow = Caution, off-nominal, trip, manual override, or attention required
  Blue = Advisory
  Green = -Not Reserved-
  Gray = Unavailable, Unselectable
Workstation Illumination
  Provide interior lighting for crew tasks and functions
Notes: Rapid Freeze = current ISS capability
Failure Tolerance = could be designed into systems or keep available spares
Configurable Setpoints = allows use of multiple temperatures simultaneously within the given ranges
Temperature Stability = JSC engineering states +/- 3C as the current fluctuation we see in sample temp stability

## Slide 44: DRAFT Human Factors Requirements (2/2)
All requirements are considered draft and subject to change prior to final RFP release.
Cleanliness
  Conform to Visibly Clean – Sensitive standards in accordance with GSDO-RQMT-1080 unless otherwise stated
Ahtropometric Dimensions Standards
  Accommodate USG crew anthropometric dimensions and ranges of motion provided in NASA-STD-3001 V2
Human Interface Usability
  Provide human interfaces that meet the intent of an average NASA-modified System Usability Scale (NMSUS) score of 85 or higher
Notes: Rapid Freeze = current ISS capability
Failure Tolerance = could be designed into systems or keep available spares
Configurable Setpoints = allows use of multiple temperatures simultaneously within the given ranges
Temperature Stability = JSC engineering states +/- 3C as the current fluctuation we see in sample temp stability

## Slide 45: Other DRAFT Applicable Requirements
All requirements are considered draft and subject to change prior to final RFP release.
Condensation Limits
  Limit condensation persistence on interior surfaces to 1 hour per day within habitable volume throughout all flight phases
Trash Management
  Contain, manage, provide odor control, and disposal for all trash, including biological waste and wet trash
Inventory Management
  Available to crew and ground and compatible with stowage system
Potable Water
  Must meet the physiochemical and microbial limits set in Appendix D.1 of CLDP-REQ-1130
Notes: Rapid Freeze = current ISS capability
Failure Tolerance = could be designed into systems or keep available spares
Configurable Setpoints = allows use of multiple temperatures simultaneously within the given ranges
Temperature Stability = JSC engineering states +/- 3C as the current fluctuation we see in sample temp stability

## Slide 46: Questions?
All requirements are considered draft and subject to change prior to final RFP release.