Hydrogen Sensor via Plasma Techniques Development*

Sarah Feldman a, Kenneth Engeling, PhD. b, Carolina Franco, PhD. b, Nicholas Spangler c, and Jessica Schwend b

- University of Michigan (sarfeldm@umich.edu)

(b) NASA Kennedy Space Center Applied Chemistry Laboratory (kenneth.engeling@nasa.gov,  carolina.franco@nasa.gov, jessica.schwend.@nasa.gov)

(c) NASA Kennedy Space Center Applied Physics Laboratory (nicholas.b.spangler@nasa.gov)

Currently, NASA Kennedy Space Center’s Exploration Ground Systems (EGS) uses liquid hydrogen (LH2) as fuel for launches and ensures hydrogen is no longer in the fill lines

Figure 1 - Simple schematic of the plasma-based hydrogen sensor setup (not to scale).

by sampling gas into a controlled environment. Then they use a catalytic sensor that detects when hydrogen interacts with oxygen. However, this procedure requires repeatedly backfilling and flushing with helium, which can be wasteful during a global helium shortage and expensive for each sampling port. Therefore, the team at KSC set out to establish proof-of-concept of a plasma-based hydrogen sensor that is anaerobic – and can in fact detect in most environments and below atmospheric pressures – and with a small footprint and more sensitive than other hydrogen sensors currently on the market.

The technology development was done by testing known concentrations of hydrogen in argon gas fed through a vacuum cube containing an electrode feedthrough at varying pressures. The resultant emission spectra were recorded with a fiber optic spectrometer and analyzed to determine the instrument's accuracy. Throughout testing, efforts were made to prove the off-the-shelf capabilities of the setup. The traditional high voltage AC-power source was switched to an affordable, handheld plasma lighter. Additionally, the spectrometer was supplemented with a double photodiode circuit to take targeted measurements of the Balmer-α and - β lines in the hydrogen spectrum.

From this, we established a proof of concept sensor. SLS required it to detect as low as 100 ppm whereas we detected hydrogen in concentrations as low as 50 ppm and in an anaerobic environment.

* Work supported by NASA Kennedy Space Center’s Science Mission Directorate Innovative Research and Development Fund.