## Slide 1: Diode Laser Hygrometer (DLH) Measurements of Water Vapor on the Convair-580 during PONEX
Glenn Diskin, Josh DiGangi
Yonghoon Choi, Mario Rana
Dave Eckberg
NASA Langley Research Center
10/28/2024

## Slide 2: Diode Laser Hygrometer (DLH)
Transceiver
Reflector
laser transmit
laser return
Open path near-IR laser spectrometer measures water vapor by differential absorption over an external path
Line-locked wavelength modulation spectroscopy @ 1.4 μm
Location of DLH external path (red) on NASA P-3B.
Retroreflector on tail
“Tophat” transceiver on P-3B zenith port
Three absorption line with different line strengths
  Large measurement dynamic range
    < 1 ppmv to >4%
  Automatic line changes to optimize sensitivity, depending on conditions
Different versions for > 10 aircraft over >25 years
Instrument performance
  Uncertainty: 5% (H2O(v)), 10-15% (RH)
  Precision: 0.1% or 0.05 ppmv
  Data rate: 1 Hz, 20 Hz, others upon request

## Slide 3: DLH on the Convair-580 for PONEX
Optical Path ~13m RT
Installation should be similar to that on the NASA P-3
  Likely can modify transceiver mounting hardware rather than building new
Sensitivity at low signal is proportional to pathlength
  Path on CV-580 is about 72% of that on P-3
  But the path on CV-580 is the same or longer than that on WB-57 during ACCLIP or on Global Hawk during ATTREX
  The 13m RT path on the CV-580 is easily sufficient to meet the measurement needs of PONEX
Expect to meet with CNRC and ECCC personnel in December to discuss details of DLH installation and operation

## Slide 4: Diode Laser Hygrometer – Short Path (DLH-SP)
Background: DLH relies on a clear view between the optical transceiver and the retroreflector.
  When cloud optical density is too high, DLH is unable to make a measurement
  Dense liquid clouds can contaminate signal with backscatter, rendering optical path uncertain
  During icing conditions, the DLH retroreflector on the P-3 can become coated with ice, resulting in very low or no signal return, until the ice sublimates
New design combines traditional “long path” DLH with a shorter path between two fins onto a DC-8 window
  On the P-3, only the short path is used
Same fundamental method as DLH
Advantages of shorter path:
  Essentially complete data coverage, even during icing or high OD cloud conditions
  More localized measurements
  Insignificant cloud scatter interference
Disadvantages of shorter path:
  Shorter path can result in lower precision
  Greater chance of interference from cabin leaks, other fuselage effects